## Slide 1: 1.3 GHz Phase Averaging Reference Line
Ed Cullerton, Brian Chase
Fermilab, Batavia, IL
10/18/2001

## Slide 2: Design Concept
Original design concept from J. Frisch, D. Brown, and E. Cisneros in their paper titled, “Performance of the Prototype NLC RF Distribution System”

## Slide 3: 1.3 GHz Phase Averaging Reference Line for Fermilab SRF Beam Test Facility
Notes: Lab measurements done with a scaled version of the reference line. (long 7/8” inch RFS cable replaced with short 0.141 cable with similar temperature characteristics)

## Slide 4: Design Limitations
The performance of the reference line will be limited by two key factors:
Impedance mismatches between the directional coupler and the coaxial line.
Phase drifts in the combiner circuit.
*Careful selection of each part must be made!

## Slide 5: Design Goals and Procedure
The design goal is to minimize phase drift between taps vs temperature.
A design procedure has been developed that minimizes the phase drift between taps of the reference line by adjusting the cable length between directional couplers so that the averaged signals are on a peak of the phase ripple caused by reflections and impedance mismatches.
The phase drift due to the combiner circuit can be minimized by good part selection and proper arrangement of each component.

## Slide 6: Directional Coupler Selection
Narda 3022 20 dB directional coupler
VSWR < 1.05
Directivity > 35 dB
Typical Narda 3022 Return Loss

## Slide 7: Cable Selection
RFS LCF78-50J 7/8” Coaxial cable

## Slide 8: Design Procedure Simulation

## Slide 9: S-parameter Results of Simulation
Broadband phase – the phase is flattened at 1.3 GHz by adjusting the time delay on the forward and reverse measurement ports.
Narrow band phase – observe that the peak of the averaged phase ripple is not at 1.3 GHz.

## Slide 10: Cable Adjustment Result
The averaged phase is on the peak at 1.3 GHz. This will minimize the phase error for small changes in length of the cable.

## Slide 11: Design Procedure Continued
Adjust the cable length between coupler 2 and 3

## Slide 12: Design Procedure Simulation
The procedure is continued for each coupler in the reference line.

## Slide 13: Complete Reference Line Simulation with Adjusted Cables

## Slide 14: Simulation Results
Phase (degrees)

## Slide 15: Combiner Circuit

## Slide 16: Combiner Circuit Coax Cable Selection
Times Microwave Phasetrack 210 coax cable

## Slide 17: Mini-Circuits ZFSC-2-5-N Power Combiner Measurements
Network Analyzer phase measurement of the two outputs of the combiner over a 25 degrees C change.
~30 mDeg Separation

## Slide 18: Mini-Circuits UNAT-3 Attenuator Measurement
Network Analyzer phase measurement of the fixed attenuator over a 25 degrees C change.
150 mDeg drift

## Slide 19: ARRA N2-3854-10 Variable Attenuator Measurement
~170 mDeg drift
Network Analyzer phase measurement of variable attenuator over a 25 degrees C change. Attenuator set to 0.5 dB.

## Slide 20: Complete Combiner Circuit Measurement
10 mDeg of separation
Network Analyzer phase measurement of combiner circuit over a 25 degrees C change.

## Slide 21: Lab Scaled Reference Line
The 7/8” RFS cable is replaced by Belden 0.141 tin coated coax. The electrical length change for 1 meter of the 0.141 cable is similar to that of 30 m of 7/8’ RFS cable.

## Slide 22: Lab Scaled Reference Line Cable Length Measurement and Adjustments
Forward and reverse phase
Averaged phase

## Slide 23: Remaining Cable Length Measurements
Coupler 3 cable measurement
Coupler 4 cable measurement
Coupler 5 cable measurement
Coupler 6 cable measurement

## Slide 24: Scaled Reference Line Lab Results
~30 mDeg of phase drift between tap 1 and tap 6

## Slide 25: Conclusions
A promising design has been proposed and analyzed for the 1.3 GHz reference line for the Fermilab SRF Beam Test Facility.
Simulations and the scaled lab model have been shown to have good phase stability over a broad temperature range.