## Slide 1: Characterisation of the Medipix3: A single-photon-counting hybrid detector
DESY Summer Students Programme 2011

## Slide 2: Contents
Introduction
Detector structure
Pixel electronics
Normal pixel behaviour
Faulty pixels
Temperature influence
Conclusion

## Slide 3: Introduction
Semiconductor photon detectors:
  Huge advances in micro-electronics industry (Moore’s law)
  Very fine structures => high spatial resolution
  Popular design: the CCD (in your digital camera)
Hybrid detectors:
  Developed at CERN for particle physics
  Also very useful in photon science (X-ray research) because of the single photon counting nature
  Advantages over CCD: e.g. energy threshold
  Medipix3

## Slide 4: Detector structure
Hybrid design: two layers
Top layer:
  Actual detector medium; converts photons to electrical current
  N-type semiconductor (e.g. Si) with array of P-type implants
  Reverse bias: normally no current (except leakage)
  Photo-electric effect: electron-hole pairs
Connection:
  Bump bonds (solder)
Bottom layer:
  Signal processing electronics (actual Medipix3 chip)

## Slide 5: Detector structure
Pixel array
One pixel
Top layer
Bump bonds
Bottom layer

## Slide 6: Pixel electronics
Count photons, i.e. pulses of electric current induced in PN-junction of the top layer
Amplify & “shape” each pulse
Then compare it to a threshold (~ minimum photon energy)

## Slide 7: Setup: Hardware/software
Detector: single Medipix3 chip 256x256 pixels, 55 μm pitch
Software: Pixelman v2.1.0, July 4th 2011
        IDL 8
USB interface
(Test pulses)

## Slide 8: New developments
Tile several chips together => increase resolution
Chip mounting & cooling: design consideration

## Slide 9: Central question
How is the performance of the Medipix3 influenced by various conditions?
  DAC settings
  Temperature
Accurate comparison: use test pulses instead of actual photons

## Slide 10: Normal pixel behaviour
Global threshold DAC (digital-to-analog converter)
Threshold very low: noise
Threshold very high: no counts
In between: correct number of counts
100 test pulses (5 ke- charge)

## Slide 11: Faulty pixels
Some pixels don’t work this nicely…
  No counts (dead pixel)
  Wrong count (noise)
  Wrong count (misadjusted pixel)
a
b
c

## Slide 12
Faulty pixel count
THL = 60
Optimal THL for each temperature
Border pixels dead
Temperature influence

## Slide 13
Example images (THL = 60)
-20 °C
+20 °C
Temperature influence

## Slide 14: Why?
Default code + 20 steps @ -20 °C
Default code + 10 steps @ -20 °C
Default code @ -20 °C
Default code - 10 steps @ -20 °C
Default code - 20 steps @ -20 °C
Temperature dependency of the DAC voltages

## Slide 15: Why?
Leakage current increases with temperature
Higher near the border (crystal defects)

## Slide 16: Conclusion
Ideal: -20 °C ambient temperature (actual chip temp. ≈ +40 °C)
Even then, still some malfunctioning pixels