## Slide 1: MTCA.4 Based Local Oscillator and Clock Generation Module for the European XFEL.
Uroš Mavrič on Behalf of the MSK Group / DESY and ISE / Technical University Warsaw,
Tony Rohlev (TSR Engineering)
DESY, Zeuthen, 03.03.2015

## Slide 2: Motivation.
The RF field detection scheme for the XFEL low-level RF system uses the down-conversion of 1.3 GHz pulsed RF. The frequency translation is performed through a mixer which mixes the LO (1.354 GHz) and the RF (1.3 GHz) down to IF (54 MHz).
The IF (54 MHz) is sampled in a fast ADC with Fs = 81.25 MHZ.
RF
LO
IF
25 RF stations
50 MTCA.4 Crates (master and slave)
9 LO and 9 CLK tap points per crate
450 LO tap-points and 450 CLK tap points
Front-panel
RF Backplane

## Slide 3: Specifications for LO and CLK Signals.
Specs for additive noise are derived from the noise contributions of other subsystems (RF front –end and ADC).
2 fs (1 mdeg at 1.3 GHz ) [10 Hz - 10 MHz]
Long term stability
0.2 ps_pp [forever – 10Hz]
Isolation on all ports
< -80 dBc
CLK additive noise
< -160 dBc/Hz floor for white noise
LO Output : 1 W at 1.354 GHz
Intermediate frequency [10MHz, 50MHz]:
CLK
BPF
DDC
CIC Filter
Calibration
Front-End <-150dBc/Hz
LNA
LO and CLK Generation
Reference
ADC
Single Channel Receiver
ADC+ DDC -147dBc/Hz
-150dBc/Hz
AM
PM
Mixer:
LO:
-150dBc/Hz – 10log10(8) = -160dBc/Hz
Vector sum processing gain

## Slide 4: Cable vs. Backplane LO and CLK Distribution.
External LO and CLK modules allow for better performance of the generated signals.
Better Temperature and humidity control of the distribution system (cables).
No external cables needed
Compact system
Ext. LO and CLK
Integrated LO and CLK
generation
Introduction of the new concept of integrating the LO and CLK distribution into the MTCA.4 crate.

## Slide 5: uRF-Backplane.
The uRF-backplane is used for the distribution of the LO, CLK and calibration signals.
The uRF-Backplane includes also other comm. signals such as I2C to modules, power lines, dedicated LVDS conn. etc.
From the point of view of MTCA.4 management it can be treated as an extension of the front AMC backplane.
It can accept 2 power modules that comply with MTCA.4 power specifications.
It can accept 4 eRTM modules.
DRTM-LOG1300 is an eRTM sitting in slot 15 in the rear.

## Slide 6: DeRTM-LOG1300.
Double-width, double full-size (12HP) module
Composed of several sub-modules
  RF daughter card
  Carrier mezz. + RF distribution mezz.
  DC/DC mezz.
  TEC mezz.

## Slide 7: DeRTM-LOG1300 - RF.
Need to use small surface mount components because of compactness -> performance is deteriorated.
LO generated from REF via dividers
Features:
  Variable LO output power by -3dB
  High resolution temperature sensor (24 bit ADC with NTC Thermistor)
  Dividers in the range from 1 to 64
  2 variants (IF=54 MHz, IF=36 MHz) – assembly defined
  Can cover RF/LO frequency range from 720 MHz To 3 GHz – assembly variant
  3 Peltier modules for temperature regulation
  All voltages on all chips are monitored
Heat-sink
Peltier Modules
Insulation - Top
RF Shield - Top
PCB
RF Shield - Bottom
Insulation - Bottom

## Slide 8: DeRTM-LOG1300 - Carrier.
Module that splits the RF, CLK signals and interconnects to the uRF-backplane.
Features:
  9 x LO/REF/Pilot outputs on Radiall connectors
  22 Diff. LVPECL CLK outputs on ERNI connectors
  Switching OFF/ON each individual CLK, LO, REF and Pilot Output
  Monitoring of the main voltages and currents
  Temperature and humidity measurements
  MMC 1.0 compliant
  Application microcontroller
  Connectivity to the ext. world
Heat-sink
Insulation - Top
PCB

## Slide 9: Other Mezzanine Modules.
DC/DC power converter mezzanine
  +12V into +5.9V and +5.4 V
RF splitting mezzanines
  For splitting the REF, LO and calibration signal
Temperature controller mezzanine
  Integrated 3 temperature controllers
Temperature Controllers
RF Input
9 x RF Outputs
RF switches
DC/DC Converters

## Slide 10: Res. Phase Noise Analysis.
2 DUT measurements of residual phase noise of individual subsystems.
Test setup 1.3 fs [10 Hz – 10 MHz]
Main limitations:
  Dividers – 3.2 fs
  Pre-amp – 3.4 fs
Overall system res. jitter = 4.3 fs (2.7 fs with de-embedded setup)
[10 Hz – 10 MHz]
1.3 fs [10 Hz-10 MHz]
3.4 fs [10 Hz-10 MHz]
3.2 fs [10 Hz-10 MHz]
4.3 fs [10 Hz-10 MHz]
Overall Res. Phase Noise
Other spurs:
54 MHz - unshielded setup
1 kHz - unknown
Other spurs come from the test setup

## Slide 11: Measurements I.
RF daughter board:
  Isolation between channels
  Return loss at connector
  Harmonic content in output signals
  Output power
The RF daughter card consumes 11 W
Reflection at [GHz]: | S11 [dB] – Measured-Shield
Ref In (1.3) | -24
Ref Aux Out (1.3) | -27
Ref Out (1.3) | -26
Cal Out (1.3) | -29
LO Out (1.354) | -26
CLK Out (1.3) | -23
LO Mon Out (1.354) | -32
CAL In (1.3) | -29
Power out | Power [dBm] - Expected | Power [dBm] - Measured
Ref Aux Out (1.3) | 26.2 | 25.8
Ref Out (1.3) | 13 | 14.7
LO Out (1.354) | 31.0 | 30.2
CLK Out (1.3) | 9.2 | 10.0
LO Mon Out (1.354) | 14 | 16.5
Power out | Shielded [dBc]
Ref Aux Out (1.3) | < -80
Ref Out (1.3) | < -80
LO Out (1.354) | < -80
CLK Out (1.3) | < -80
LO Mon Out (1.354) | < -80
Pilot | < -80
Power out | 2nd [dBc] | 3rd [dBc]
Ref Aux Out (1.3) | <-80 | <-80
LO Out (1.354) | <-80 | <-80
CLK Out (1.3) | <-80 | <-80
REF (1.3) | <-80 | <-80
Return Loss
Isolation between Ch.
Output Power
Harmonic Content

## Slide 12: Measurements II.
S parameters of the splitting section:
  S21 = LO -16 dB (spread = 0.4 dB), CAL -16 dB (spread = 0.5 dB)
  S11 = < -22 dB
  Isolation = mostly < -80 dB, some specific channels -65 dB
LO and CLK Distribution over the uRF-Backplane:
Amp
Pha
Sampling of signals with CLKs that were distributed over the uRF-backplane. No additional spurs were visible.
CLK distributed over the RF backplane
to slot 4 (long. distance).

## Slide 13: System Integration.
Subsystems Involved:
  DRTM-LOG1300
  uRF-Backplane
  9U Chassis
  NAT-MCH-BM or Rear Power module
  End-Users (RTMs)

## Slide 14
Thank you for your attention!