## Slide 1: WG2; Lab Contributions about A) Instrumentation; Limits; Procedures incl. Discussion
Fumio Furuta
Cornell University
TTC Meeting WG2, 3Dec2015, SLAC

## Slide 2
VT area | probe | Interlocked | location | Notes
Inside of rad. shielding block | 3 Gamma | No | Semi-fixed in pit | 
Outside of rad. shielding block | 3 Gamma | Yes | Fixed on wall | 2mR/h limit.
 | 2 Gamma | No | Flexible | 
MLC | Probe | Interlocked | location | Notes
Inside of rad. shielding block | 3 Gamma | No | Semi-fixed on MLC (coupler, end plates) | 
Outside of rad. shielding block | 3 Neutron | No | Fixed on wall | 
 | 6 Gamma | Yes | Fixed on wall | 2mR/h limit.
 | 1 Gamma | No | Handy for survey | 
Cavity VT area at Newman lab basement.
MLC test area at Wilson L0E.
1) Which diagnostic techniques wrt. radiation are applied? -> gamma proves

## Slide 3
Cavity VT area in Newman lab. basement

## Slide 4
MLC String and Rad monitors
Gamma probe #1, #2, and #3 on MLC.
#2 was put on the coupler of measured cavity.
#1
#2
#3

## Slide 5
Wilson L0E for MLC test

## Slide 6
FE onset
FE onset
Cavities were counted as FE limited when
Qo dropped with rad increasing,
Showed high Rad (>1~10R/h) w/, or w/o, quench.
2) What is your criteria for "field emission" in the test?
What is your criteria for counting a cavity as FE loaded/limited/relevant?
VTs limited by FE

## Slide 7
3) What is your guideline for the vertical test in case of field emission present?
If cavity achieved design specs, stop as admin. limit.
If not, keep rising RF power .
If Cavity had quench, try processing by keep quenching,,
in Pi-mode (single-cell, multi-cell)
in Pass-band (multi-cell)
Abort test when
input power was too high, need to prevent burn RF cable.
cavity heating was too high to maintain temp.
Rad at outside shielding > 2mR/h, safety limit.

## Slide 8
4a) Which processing techniques are applied? -> quenching.
What is the success rate? -> table
4b) How often is a spontaneous degradation observed (Q-value, radiation, gradient)? -> table
 | VT | HTC7-cell | HTC-9cell | MLC
# of test | 53 | 3 | 5 | 6
# of FE limited | 10 | 1 | 1 | 2
Eacc degradation | 2 | 1 | 0 | 1
Qo degradation | 8 | 1 | 1 | 2
Applied processing | 10 | 0 | 1 | 2
success | 0* | 0* | 1 | 1
Total success rate [%] | ~15% (2/13) |  |  | 
* Some improvements were seen, but not fully processed out

## Slide 9
5) Percentage of cavities showing field emission? 6) Reprocessing strategy?
 | # of Cavity | # of VT | # of FE | cured by | 
 |  |  |  | re-HPR | re-chem.
total | 17 | 53 | 10 | 3 | 7
Chart: 
 | 43 | 3 | 7
FE limit ~19% of VT (multi-cells)
Reprocessing strategy
Optical Inspection
based on T-map, OST data if available.
Reprocess will depend on R&D goal, schedules, or budget.
Re-HPR ; if schedules were not tight
Re-Chemistry ; BCP or VEP, depends on R&D target.
Re-HPR
Re-chem.
FE limited

## Slide 10
7) If available: Comparability between horizontal/module and vertical test wrt. radiation/dark current?
Ratio of test limited by FE (# of FE limited / # of toal tests) |  | 
VT | HTC | MLC
15% (5/33) | 25% (2/8)* | 3% (2/6)**
FE limit during VT / HTC
*, ** One cavity processed out by quenching

## Slide 11
Example of FE processed out
Quenching processed out FE.
1st process
2nd process

## Slide 12
8) is there any relation between Multipacting and FE observed
Chart: #
 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
 | 1 | 1 | 1 | 1
 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
 | 1 | 1 | 1 | 1 | 1 | 1 | 1
 | 1 | 1 | 1
Chart: Eacc [MV/m]
 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 2 | 0 | 0 | 1 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Histogram of gradient with FE, analysis on 11 test limited by FE.
Histogram of FE onset, analysis on 11 test limited by FE.