## Slide 1: Lessons learned from cryomodule assembly toward future projects
TTC 2015| Catherine MADEC
other talks on WG3 By T. Trublet and S. Berry, WG4 by H. Jenhani

## Slide 2: Cryomodules assembly at CEA
Booster
MACSE
SOLEIL
TTF
SPL
SPIRAL 2
ESS,
SARAF,
IFMIF
XFEL
l/4
E
B
100
12
1
3(+1)
30

## Slide 3: SRF dedicated INfrastructures
CourtYard
©GoogleMap
Clean rooms
Assembly halls
Labs
Warehouse
Chemistry
Cryo-RF area
Notes: This is an sky view of the building dedicated to XFEL project in Saclay
The Xfel village is made of clean rooms, assembly halls, officces and warehouse

## Slide 4: XFEL Workstations
Coupler Area CO-WS1 & 2
Alignment Area AL-WS1 & 2
Shipment Area
SH-WS1 & 2
Roll-out Area
RO-WS1 & 2
Warehouse
Clean room Area
CC-WS1
SA-WS1 & 2
Reception Area
REC-WS1b
Cantilever Area CA-WS1
Reception Area
REC-WS1a
SRF facilities
SRF facilities
Assembly area
Chemistry
Notes: This is the top view scheme of the XFEL village.
One can see the different workstations in the assembly halls.
The workflow is the following
This is the warehouse ;
Parts have to pass through reception tests on this area.
The clean assembly of the most critical compounds occurs in a 100m2 ISO4 clean room.
Pursue assembly by rolling out the cavity string from the CR
The Alignment area is in the same hall
/cross the yard with special truck/
A big cantilever helps to hang the cold mass while sliding the vacuum vessel around it
This workstation is mostly dedicated to warm part of the coupler and its pumping line.
We have to move in another hall to check and prepare the module for shipment.
The RO area is the workstation where the cavity string is dressed with ancillaries
The

## Slide 5: Project phases : ALL MODEL works even if not present at all the steps but
QC easier when at all steps including design
especially when one has to fix a problem
everything is in the details
If at design, the details are known, other problems will show up later but you’ll be able to handle
For ESS, CEA has to set up parts manufacturing follow up
 | SPIRAL2 | XFEL | ESS | SARAF | IFMIF
Design | X |  | X | X | X
Preliminary study | X | X | X | X | X
Preparation of Infrastructure and tooling | X | X | X | X | x
Prototypes manufacturing | X |  | X | X | X
Mock-ups (training, tooling rework, assy procedure tests) | X | X | X | X | X
Prototyping | X | X | X | X | x
Parts manufacturing | X |  | X | X | X
Chemistry | X |  |  | X | X
Assembly | X | X | X | X | x
Test en bunker | X |  | X | X | 
Transport | X |  |  | X | x indiv
X : CEA has been involved
X : not present at this step
where not present
where not present

## Slide 6: Assembly models at CEA
The assembly were performed :
CEA ressources for SPIRAL2
Subcontracted for XFEL
Main differences :
The thoughput was different (1CM/wk for XFEL),
The number of modules produced (100/12)
Lessons – learned :
Robust assembly procedures (which are not sensitive to the environment – SPIRAL 2 : 5 different assembly location 3CR @Saclay; IPNO, CERN )
Adapted to project
      - Analysis (drawback and advantages) of each scenario (industrial or CEA ressources) for each project
SP2 Coupler-cavity assembly in 2 different CR at CEA

## Slide 7: Lessons learned on quality control
    Random checks at manufacturer
    Inspection at CEA on critical components (cavities-valves-bellows = beam vacuum components) otherwise when ready to assemble
    Assembly late after SAT (site acceptance test – no warranty) rework by CEA
Pour insérer une image :
Menu « Insertion / Image »
ou
Cliquer sur l’icône de la zone image
 | SPIRAL2 | XFEL
Inspection of incoming parts | x | x
Controlling assembly | Auto | X
Documentation | X | X
Audits | x | x
X
Auto
X
    Autocontrol by the technician themselves
    Improvment of the assembly procedure along the assembly
    Documentation : Record of all the steps, tests,
x
    Audit on coupler -> improvement
X : random or rarely
X: on everything

## Slide 8: Lessons learned list on quality control
 | SPIRAL2 | XFEL
Inspection of incoming parts | X | x
Controlling assembly | Auto | X
Documentation | X | X
Audits | x | x
Pour insérer une image :
Menu « Insertion / Image »
ou
Cliquer sur l’icône de la zone image
x
    Manufacturer qualification by our lab partners
    -> parts come at CEA conformed
    Inspected at CEA :
    when possible at arrival
    not everything can be inspected
    so during assembly
    Stock so NCR parked (and repaired later) and
    replaced by next part
    NCR : parts reworked, sent back or trashed
X : random or rarely
X: on everything

## Slide 9: Lessons learned list on quality control
 | SPIRAL2 | XFEL
Inspection of incoming parts | x | x
Controling assembly | Auto | X
Documentation | X | X
Audits | x | x
    Until end of 2014, the quality control group of Alsyom was too small (3 people) so they performed Incoming Inspection and Documenting in priority. CEA took over the Assembly Work Controling (‘Hold Points’). The ‘every day’ or ‘random’ controls were too few and this led to many mal-fabrication, most of them recorded at DESY before or during cryomodule cold test
    In November 2014, Increase of the quality control group of Alsyom to 5 people and a better organisation, cover the need of QC. In August 2015, increase of CEA quality control group.
    CEA also took over the NCR editing on EDMS and improvment proposal
    Audits (XM26 + local ones) are good – Feedback from the auditors (CEA experts) to the operators
X
X
X
x
X : random or rarely
X: on everything

## Slide 10: Foreseen plan for quality control
 | SPIRAL2 | XFEL | ESS | SARAF | IFMIF
Inspection of incoming parts | x | x | x | x | x
Controling assembly | Auto | X | X | Auto | x
Documentation | X | X | X | X | X
Audits | x | x | x | x | x
    ESS : XFEL like with CEA efforts in order to check the components at manufacturer and at arrival (CEA ressources or subcontractor) and during the assembly
    SARAF : SPIRAL2 like
    IFMIF : new paradigm because components will be inspected and prepared at CEA control and shipped individually to Rokkasho for assembly
Pour insérer une image :
Menu « Insertion / Image »
ou
Cliquer sur l’icône de la zone image
X : random or rarely X: on everything

## Slide 11: Lessons learned : industrial contract
Chart: CRYOMODULE NUMBER

XM46
XM50
XM21
XM1 March 2014
5 day throughput was reached mid-October 2014 with XM15
4-day throughput was reached in January 2015 with XM25
Acceleration needed to close tunnel mid-2016
XM77 to be delivered at the end of December 2015
      Contract signature when everything thought off – low engineering input
      No industrial was trained for cryomodule assembly
        CEA did the training of 8 first people and more 2/3 (good but CEA ressources - experts)
        The detailed industrial analysis lead to re-optimisation of the work flow in autumn 2014
      Infrastructure set up to the right level. We can accomodate the less than a wk throughput and rework on modules (XM46&XM50) ; Detailed by S. Berry and T. Trublet in WG3
7wks
6wks

## Slide 12: Lessons learnED : improvment
  Output of the prototyping phase, the assy procedure written in english and appended to the call for tender. Alsyom translated in french and updated. This took much longer than expected Now on, in FRENCH then translated if needed/required
  Aging of tooling on XFEL (starting around 20th module) – clean room tooling spares
  Propose new solutions at prototype stage :
    MLI
    Magnetic shield
  Implement new tooling meanwhile producing
    Socks on cavity posts
    gate valve assembly in CC area
    leak detection
Pour insérer une image :
Menu « Insertion / Image »
ou
Cliquer sur l’icône de la zone image
    Socks used to transfer cavity string from Clean Room to Rollout Area, reduce transfer time (no cleaning)
Realization of gate valve support for its assembly on Cold-Coupler workstation
Leak detection by accumulation

## Slide 13: improvment on assembly which could IMpact performance
New clean room assembly procedure: moving individual cavity venting after the leak check of the cold coupler assembly, rather than later before the string assembly. This eliminates one connection to pumping stations for cavity venting, and one valve closing-opening cycle. Cavity equipped with filters to be connected to N2.
Tested on XM27 and implemented from XM54
Pour insérer une image :
Menu « Insertion / Image »
ou
Cliquer sur l’icône de la zone image
New clean room procedure
Chart: Average RF gradient (MV/m) for taylored RF distribution/individual cavities

## Slide 14: Impacted performances
Eacc Spec = 23,6 MV/m Eacc Max = 31 MV/m
Average gradient is 16% above specs : Eacc = 27.5 MV/m.
XM6 to XM23, degradation while CEA and Alsyom put all their effort in achieving production goal of 1 CM/week
XM26 : audit
XM45: Electrical shutdown in clean room during assembly while the string of 3 cavities was under N2-flushing
XM46: Leak on the beam line in CR repaired, 2nd leak at the end of the assembly – Repaired at Saclay, 2 extra venting-pumping cycles
Chart: Average RF gradient (MV/m) for taylored RF distribution/individual cavities
XM45
XM46
Clipping at 31 MV/m is done due to max. available RF power; limit given by waveguide distribution.
XM26
-2.1MV/m
-0.9MV/m
0.4 MV/m

## Slide 15
SPIRAL 2 Cryomodules performances
Courtesy of C. Marchand, SRF 2015, Whistler
Q0 |  | CMA | CMB
Specifications (10 W limit) |  | 4,0E+08 | 1,4E+09
Computed (simulation codes) |  | 7,6E+08 | 2,7E+09
Tests in VC @ 6.5 MV/m Deduced from RF losses | mean | 5,8E+08 | 3,7E+09
Tests in CM @ 6.5 MV/m Deduced from therm. losses | mean | 5,0E+08 | 3,0E+09
Vertical cryostat: Pcav = PRF
Cryomodule: Pcav = Ptot – Pstat
measured via thermal losses @ 4K
(A: He level decrease; B: gas flow)

## Slide 16
      Cryogenic fluid inside the IFMIF cryomodule  Must comply with High Pressure Gas Safety Law
      Strategy negotiated with the Japanese authorities: design, fabrication and tests of all the cryomodule components according to ASME BPVC or B31.3
      HWR is the most complicated part to be licensed due to the use of non referenced materials (Nb, NbTi) and complicated geometry
The licensing procedure with the Japanese Authorities started with this component in Dec 2013.
7 meetings with KHK were necessary prior to the official submission in June 2015.
The licensing activity has been an unexpectedly time and resource consuming activity to perform the numerical simulations, deal with the many questions raised by KHK, prepare and test the samples (NbTi, Nb, NbTi-Ti, NbTi-Nb, Nb-Nb), and complete the application form.
Licensing of niobium cavities
For SARAF,
We agreed on
PED 3.3

## Slide 17
Needles magnetized
Bearing cage
Brass
cage
Ceramic needles
Homemade rollers
Off-the-shelf rollers
Homemade rollers
  Mitigation
  Mitigation
      Use non ferromagnetic materials
  Magnetization hazard
        Solenoid packages and cryogenic stepper motors inside the magnetic shield
       Risk of magnetization of some components by the fringe field of the solenoids during operation of the cryomodule
      Ex: Off-the-shelf bearing originally intended to be used
Lessons learned : Mitigation to minimize magnetization issues
Helps SARAF design

## Slide 18
      New support frame in Ti40 instead of stainless steel 316L devised
      No more pb with the welds
      Additional benefits: shrinkage lowered, cooling time divided by 3
      Supporting elements of the phase separator in Ti40
  Studies and measurements of all parts inside the
  magnetic shield of the cryomodule liable to be magnetized or naturally magnetized
  Motors, rollers, circlips, Invar tie rod, rolling components and c-shape elements
  supporting the cavities, bearings of the tuning system, etc.
  Cf. PLOUIN et al, “CEA experience and effort to limit magnetic flux trapping in
  superconducting cavities”, SRF2015, TUPB100
  Procurement specifications and control procedures
  Systematic checking of the components liable to be magnetized before cryomodule assembly
  Additional studies to be performed
  Work out degaussing procedures to eliminate any magnetic pollution
  cf. R. LAXDAL, Review of the Magnetic Shielding Design of Low-Beta Cryomodules, SRF2013, WEIOD01
Mitigation to minimize magnetization issues

## Slide 19
  Manufacturing phase ongoing
      Contract launched in Dec. 2014
      Qualification of the welds according to ASME BPVC ongoing (with a Lloyd’s register third part inspector)
      Pre-serieries cavity to be delivered at Saclay in Dec. 2015 for chemical etching
      cf. G. Devanz et al, Progress in IFMIF half wave resonators manufacturing and test preparation, THPB045
      First test in SaTHoRI test bench by early 2016
Manufacturing status
cavity torus, central drift tube, conical stem
Feedback to SARAF CEA team

## Slide 20
Manufacturing status of the other main Cryomodule components
  Manufacturers selected or calls for tender on going
        Phase separator, Support frame, magnetic shield
  Technical specifications being prepared
        Current leads and solenoid packages (CIEMAT),
        MLI, cryo and vacuum piping (CEA)
  Ongoing fabrication of key cryomodule components
                Power couplers
                Vacuum tank, thermal shield
                tuning systems
Talk by H. Jenhani talk on WG4a

## Slide 21: ESS Cryomodule design
Similar to CEBAF/SNS cryomodule concept with 4 cavities per cryomodule
Common design for medium (6 cells) and high beta (5 cells) cavities
Proton Beam
Jumper connection
Cold to warm transition
Magnetic shielding
Power coupler
Cavity with Helium tank
Diphasic He pipe
Thermal shielding
Regulation He valve
Intercavities belows
Vacuum valve
Spaceframe support
Heat exchanger
Alignement fiducial

## Slide 22
Spaceframe
Cavity supporting rods
Three supports at 120°C for the spaceframe
Access trap
Cold tuning system
Bursting disk
Cryomodule support (no ground support design yet)
Vacuum vessel

## Slide 23: ESS : Looking at the contrat model
  Time to reach nominal throughput longer than excepted (ESS – XFEL)
ESS number of modules
ESS : evaluating the best model
for the production – throughput 1 cryomodule /
3 wk

## Slide 24
ESS prototyping phase : transfert
H
124E
126

## Slide 25: ESS prototyping phase : clean room assy
124E clean room (CR)
ESS cavity string in CR
on carriage
Views of ESS cavity string in CR
on carriage

## Slide 26
ESS prototyping phase : transfert
H
124E
126

## Slide 27: Both for ESS CRYOmodule assembly

## Slide 28: Both for ESS CRYOmodule assembly

## Slide 29
ESS prototype Cryomodule assembly

## Slide 30
ESS prototype Cryomodule assembly

## Slide 31
ESS prototype Cryomodule assembly

## Slide 32: SARAF cryomodule TRANSPORT
| PAGE 32
Transportation
      Transportation by sea = Cheapest solution but:
      Several loadings/unloadings  heavy shocks possible during transshipment
      Coupler window failure or weakening due to fatigue (resulting from ocean swell)
      Risk of contamination of the beam vacuum during shipment (long journey)
Mitigation action:
Transportation by plane
Design complementary constraints: dimensions of the doors of the widespread 747-8F cargo plane: the nose door is 94 in. (239 cm) wide and 96 in. (244 cm) high, while the large side cargo door allows to charge 238.5 in. (606 cm) long and 96 in. (244 cm) wide loads

## Slide 33: Conclusions (1)
 | Lessons learned
Design | 
Preliminary study | 
Preparation of Infrastructure and tooling | 
Prototypes manufacturing | 
Mock-ups (training, tooling rework, assy procedure tests) | 
Prototyping | 
Parts manufacturing | 
Chemistry | 
Assembly | 
Test en bunker | 
Transport | 

## Slide 34: Conclusions (2)
With the various number of modules, we adapt to new situation
Manpower estimation especially in QA/QC should never be minimised
Infrastructure eases assembly and prototyping
Scrutinize project for other projects

## Slide 35: DSM Irfu SACM
Commissariat à l’énergie atomique et aux énergies alternatives
Centre de Saclay | 91191 Gif-sur-Yvette Cedex
T. +33 (0)1 69 08 69 39 | F. +33 (0)1 69 08 99 89
  Etablissement public à caractère industriel et commercial | RCS Paris B 775 685 019
THANK you for your attention