## Slide 1: Detector Commissioning and Calibration
Markus Kuster
for WP75
XFEL Commissioning Meeting
DESY, Hamburg
July, 15 2015

## Slide 2: Detectors for the European XFEL
Pixel
Size
500 μm
200 μm
75 μm
50 μm
30 μm
Energy [keV]
0.2
1.0
10
20
30
DSSC
AGIPD
LPD
Gotthard V2
pnCCD
MCP
FastCCD
4.5 MHz
4.5 MHz
SASE 1/2 (2.3 – 25 keV)
SASE 3 (0.25 – 3 keV)
4.5 MHz
4.5 MHz
4.5 MHz
4.5 MHz
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 3: Detector Mapping to the Scientific Instruments at the European XFEL – “Primary Customer”
SASE 1
SASE 3
High E
SASE 2
Single Particles,
Clusters and
Biomolecules (SPB)
Materials Imaging &
Dynamics (MID)
Small Quantum Systems (SQS)
Spectroscopy and Coherent
Scattering (SCS)
High Energy Density
Matter (HED)
AGIPD
Femto Second X-ray
Experiments (FXE)
Gotthard V2
Fast CCD
DSSC
Gotthard V1
LPD
MCP
DSSC
Small pnCCD
MCP
AGIPD
Gotthard V2
Gotthard V2
Gotthard V2
Fast CCD
Fast CCD
Low E
4.5 MHz
4.5 MHz
4.5 MHz
4.5 MHz
4.5 MHz
4.5 MHz
4.5 MHz
4.5 MHz
4.5 MHz
4.5 MHz
4.5 MHz
Notes: FXE:
will permit ultrafast pump-probe studies of solid and liquid samples, using different structural tools: X-ray absorption spectroscopies, requiring tunability of the incident monochromatic (E/E down to a few 10-5) radiation, X-ray emission spectroscopies (XES, RIXS, etc.), and X-ray diffraction (XRD) capabilities (WAXS, SAXS, Bragg) requiring both large (1 Mpx) and small (< 128 kpx) area detectors.
ultrafast studies of nuclear, charge, and spin dynamics; dynamics of solids and phase transitions, surface studies via X-ray reflectivity; dynamics of photoactive proteins; chemical dynamics in the solution phase including solvation dynamics
SQS:
The instrument will be suitable primarily for high-field applications, non-linear processes, and time-resolved pump-probe investigations. The instrument will be suitable primarily for high-field applications, non-linear processes, and time-resolved pump-probe investigations. For the investigation of multiple ionization and multiple-photon excitations, the fast relaxation of these core holes through Auger decay, which takes place on the femtosecond time scale, results in complex dynamics governing the sequential and direct ionization processes. Dynamics of atomic processes (e.g. inner shell dynamics).

## Slide 4: Detectors for Day One Operation at XFEL.EU
Adaptive Gain Integrating Pixel De-tector (AGIPD)
  Energy Range
  3 – 13 (25) keV
  Dynamic Range
  104 ph/px/pulse@12 keV
  Single Photon Sens.
  Yes
  Memory ≈380 images
  Pixel Size 200×200 μm2
DePFET Sensor with Signal Compres-sion (DSSC)
  Energy Range
  0.5 – 6 (25) keV
  Dynamic Range
  6000 ph/px/pulse@1 keV
  Single Photon Sens.
  Yes
  Memory ≈800 images
  Pixel Size 236×236 μm2
Large Pixel Detector (LPD)
  Energy Range
  3 – 13 (25) keV
  Dynamic Range
  105 ph/px/pulse@12 keV
  Single Photon Sens.
  Yes
  Memory ≈512 images
  Pixel Size 500×500 μm2
FastCCD
  Energy Range
  0.25 – 6 keV
  Dynamic Range
  ≈360 ph/px/pulse@1 keV
  Single Photon Sens.
  Yes
  Memory No
  Pixel Size 30×30 μm2
4.5 MHz
4.5 MHz
4.5 MHz
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 5: Detectors for Day One Operation at XFEL.EU
Adaptive Gain Integrating Pixel De-tector (AGIPD)
  Energy Range
  3 – 13 (25) keV
  Dynamic Range
  104 ph/px/pulse@12 keV
  Single Photon Sens.
  Yes
  Memory ≈380 images
  Pixel Size 200×200 μm2
DePFET Sensor with Signal Compres-sion (DSSC)
  Energy Range
  0.5 – 6 (25) keV
  Dynamic Range
  6000 ph/px/pulse@1 keV
  Single Photon Sens.
  Yes
  Memory ≈800 images
  Pixel Size 236×236 μm2
Large Pixel Detector (LPD)
  Energy Range
  3 – 13 (25) keV
  Dynamic Range
  105 ph/px/pulse@12 keV
  Single Photon Sens.
  Yes
  Memory ≈512 images
  Pixel Size 500×500 μm2
Other Detectors
  1D detectors for high repetition rate applications
  Small area, low repetition rate, low energy 2D detectors
  CCDs for low speed imaging
  Point like detectors for veto applications
5
4.5 MHz
4.5 MHz
4.5 MHz
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 6: Steps Towards First Day of Operation
Before delivery of the detector to XFEL
  Early performance testing and commissioning of prototype systems (single modules/quadrants) together with developer groups with in-house and external sources
  Control and DAQ integration as early as possible
  Participate in and perform performance optimization and calibration together with development groups
XFEL
AGIPD
DSSC
LPD
Delivery of the detector to XFEL
  Verification of detector performance, control system and DAQ/data management system “end-to-end” test in the XFEL detector laboratories at HERA South without XFEL beam
  Commissioning and initial calibration of the detector at HERA South before integration of the system into the beam line experiment
Detector Integration into and Commissioning at Beam Line
  DAQ and control system installation will take place before the detector will be installed
  Installation of the detector system and functional tests (vacuum, mechanics, control system, DAQ system, data management system) → repeat end-to-end test at beam line
  Commissioning and alignment with XFEL beam in-situ
    Detector ready for first day user operation

## Slide 7: Steps Towards First Day of Operation
Before delivery of the detector to XFEL
  Early performance testing and commissioning of prototype systems (single modules/quadrants) together with developer groups with in-house and external sources
  Control and DAQ integration as early as possible
  Participate in and perform performance optimization and calibration together with development groups
XFEL
AGIPD
DSSC
LPD
Delivery of the detector to XFEL
  Verification of detector performance, control system and DAQ/data management system “end-to-end” test in the XFEL detector laboratories at HERA South without XFEL beam
  Commissioning and initial calibration of the detector at HERA South before integration of the system into the beam line experiment
Detector Integration into and Commissioning at Beam Line
  DAQ and control system installation will take place before the detector will be installed
  Installation of the detector system and functional tests (vacuum, mechanics, control system, DAQ system, data management system) → repeat end-to-end test at beam line
  Commissioning and alignment with XFEL beam in-situ
    Detector ready for first day user operation

## Slide 8: Calibration and Commissioning Strategy
XFEL.EU Project Phase | Detector Laboratory | FELs/Synchrotrons/ Particle Accelerators |  | 
 |  | External Sources | XFEL. EU | 
 |  |  | Instr. Hutches | Dedicated beam
Start-up (before delivery) | Initial detailed calibration and characterization of detector modules | Response of the detector depending on repetition rate and intensity. |  | 
Commissioning | Initial detailed calibration and characterization of full systems | Calibration of large area 1MPix 2D cameras is NOT feasible due to infrastructure constraints | Validate laboratory calibration Characterization of detector response under XFEL.EU conditions (repetition rate and intensity) | 
User Operation | Calibration of full systems only if necessary (accident, detailed detector response investigation, etc.) |  | Scientific Experiments Part of regular data taking procedure (shall take a few minutes) Calibration/tests using in-hutch sources | Detailed detector response investigation under realistic conditions (repetition rate and intensity)

## Slide 9: Detector Calibration & Characterization Parameters
Parameter | AGIPD | DSSC | LPD | CCDs (10 Hz) | Strip Detectors | Point Detectors
Conversion gain G(cell nr, x,y) |  |  |  |  |  | 
Dark signal O(cell nr,x,y) |  |  |  |  |  | 
Noise N(cell nr, x,y) |  |  |  |  |  | 
Bad pixel B(cell nr,x,y) |  |  |  |  |  | 
Quantum efficiency QE(x,y) |  |  |  |  |  | 
Detective quantum efficiency DQE |  |  |  |  |  | 
Dynamic range DR(x,y) |  |  |  |  |  | 
Pile-up |  |  |  |  |  | 
Memory cell droop (signal losses) MD(cell nr, x,y) |  |  |  |  | Depends on the detector | 
Charge transfer inefficiency CTI(x,y) |  |  |  |  |  | 
Point spread function PSF(x,y) |  |  |  |  |  | 
Line spread function LSF(x,y) |  |  |  |  |  | 
Spectral response |  |  |  |  |  | Depends on the detector
Flat field corrections |  |  |  |  |  | 
Common mode |  |  |  |  |  | 
Event splitting |  |  |  |  |  | 
Alignment |  |  |  |  |  | 

## Slide 10: What is needed to calibrate and commission our detectors?
Required | Parameter | Comments
Internal charge injection | Gain conversion, memory cell droop | Internal source needs to be cross calibrated to real X-rays
Laboratory X-ray sources (isotopes, X-ray tubes) | Gain conversion, Detector Response Function, Flat-field, Charge Transfer Inefficiency, Quantum Efficiency Flat-field, Splitting Events – Crosstalk | Measurements which do not require high intensity and XFEL timing structure at the same time
Beam from FELs/Synchrotrons/ Particle Accelerators |  | All measurements which require high intensity and appropriated XFEL timing
XFEL Experiment Environment | Background, noise, EMC, com-patibility with other components | Tests must be performed in-situ.
Infrastructure to run detector
DAQ and Control system, data storage, data base, cooling, cabling etc.
Data analysis tools
Different kind of stimulus (X-rays, charged particles, etc..)

## Slide 11: Detector Calibration – Priorities
Parameter | Priority | Target Accuracy
Conversion gain G(x, y) | High | < Poisson statistic
Dark signal O(cell nr, x, y) | High | < detector noise
Noise N(x, y) | High | < 1%
Bad pixel B(cell nr, x, y) | High | 
Quantum efficiency QE | High | 5-10 % (simulation uncertainty)
Dynamic range DR(x, y) | High | < Poisson statistic
Memory cell droop (signal losses) MD(cell nr, x, y) | High | < detector noise
Charge transfer inefficiency CTI(x, y) | High | < 1%
Point spread function PSF(x, y) | Low | to be defined
Line spread function LSF(x, y) | Low | to be defined
Spectral response | Medium | < 10%
Flat field corrections | High | < 10%
Common mode | High | to be defined
Event splitting | High | to be defined, for low rate applications
Alignment (position calibration) | High | < pixel size
Measurement
Physics Model

## Slide 12: Laboratory X-ray Sources vs. Requirements
Requirements | Radioactive Isotopes | X-ray Source in DC mode | Pulsed X-ray source
Energy range 0.25 – 25 keV | Different isotopes/filters | Anode with different target materials/ filters | 
Pulse Duration Pulse length < 50 ns t = 220 ns (4.5 MHz repetition rate) |  |  | Pulsed electron source
Intensity Adjustable 1- 103 photons/ pixel/frame Stability better than 1% |  | High-power micro-focusing X-ray tube | 
Illumination Point-like d<= 20-50m Line-like Flat field (homogeneity > 10%) | Optics, pinholes, collimators |  | 

## Slide 13: In-House Calibration and Commissioning Tools
Phoebe
Test of small detectors and detector tiles
Radioactive isotopes e.g. 55Fe up to a few GBq
Flexible system for vacuum operation
Little/Big Amber
Operation of single detector tiles and detector modules up to ¼ Mpix
Detector operation under controlled and dry
atmosphere conditions
In combination with X-ray optics point like or parallel beam illumination is possible
¼ Mpix DAQ and control backend
PulXar
Operation of detector modules and full detec-tors up to 1 Mpix (XHQ) ¼ Mpix (HERA South)
XFEL timing and pulse structure
Full 1 Mpix DAQ and control backend
Phoebe
Little Amber
PulXar
Big Amber
Notes: The concept is to built up infrastructure for XFEL in several steps.
Problem is …
→ go through right part of the slide
What are the specifications, requirements → next slide

## Slide 14: In-House Calibration and Commissioning Tools
Pulsed e- Gun
Standard Sources
Radioactive isotopes e.g. 55Fe up to a few GBq
Low power commercial X-ray tubes
“High Intensity” Sources
Different target materials
→ different photon energies 0.2 to 16 keV
In combination with X-ray optics point like or parallel beam illumination
Up to 106 photons/s/pixel continuous
Pulsed Electron/X-ray Source
XFEL timing and pulse structure and
continuous operation
e- current up to 20 mA
Point like and small area illumination
e- and X-ray generation
Up to 1000 photons/pixel/pulse @ 4.5 MHz
Low Power
X-ray Tube
Single-Target
X-ray Tubes
Multi-Target X-ray Tube
Notes: The concept is to built up infrastructure for XFEL in several steps.
Problem is …
→ go through right part of the slide
What are the specifications, requirements → next slide

## Slide 15: Detector Laboratory in HERA South
Notes: The concept is to built up infrastructure for XFEL in several steps.
Problem is …
→ go through right part of the slide
What are the specifications, requirements → next slide

## Slide 16: Size of Detector Parameter Space
Example LPD
x 512 memory cells
x 1 million pixel
= 5 x 108 parameters
and
3 Gain Stages
x 2 Gain Settings
16 Super Modules
256 Detector Tiles
2048 ASICs
LPD 1 Mpix Detector
Parameter Dependence
Temperature
Integration time/sampling speed
Irradiated dose
Conclusions
Impossible to calibrate and commis-sion many (>3 - 4) combinations for first day of operation (finite time)
Focus on most important operating conditions!

## Slide 17: Scientific Requirements for Day-1 Operation
Non Linear System Response
 | DSSC | AGIPD | LPD
Beam structure | Accelerator freq: ~100 kHz, 60 pulses per bunch, 600 pulses/sec, 10 μs pulse separation |  | 
Photon energy | 0.5, 0.7, 1, 1.5 keV | 7, 8.3, 9, (15) keV | 12, 15, 20 keV
Expected max. nr of ph/pulse/pixel | 3 x 103 @1keV (maximum achievable) | 104 @ 12keV(SPB) | >104 @12 keV
Dynamic range as a function of the radial detector position | YES, sensor by sensor (ASIC unit) |  | 
Single photon detection | Not requested | Yes (MID) | Yes, at 20keV (minimum requirements)
Alignment precision | Alignment on quadrant level needs to be stable and known on level of 1/10 of the pixel size |  | 
Veto capability | Yes | Yes | Yes
Simulation data (typical patterns) | Simulations of FeRh LCLS experimental data to see performance to be provided | Simulated data, initially to investigate interoperation of detector-sim with SPB s2e-sim | FXE provided simulated data to compare updated LPD specs with requirements
Other requirements | Calibration with filters |  | 

## Slide 18: Detector Operation Modes – First Day
Non Linear System Response
DSSC
 | Mode I | Mode II | Mode III
Read Out Frequency | 4.5 MHz |  | 
Photon Energy | 1 keV | 1.5 keV | 0.7 keV
Max # of ph/pulse/pixel | 3 x 103 @1keV (maximum achievable) |  | 
Read Out Geometry | Full Frame |  | 
Single Photon Sensitivity | No |  | 
Alignment Precision | 1/10 of the pixel size (quadrant level) |  | 
Veto Capability | Yes& |  | 
Memory | 800 Cells (full memory) |  | 
AGIPD
Mode I | Mode II | Mode III
4.5 MHz |  | 
8.3 keV# | 15 keV | 7 keV
104 @12.4 keV (maximum achievable) |  | 
Full Frame |  | 
Yes 0.15 ph RMS Noise |  | No
1/10 of the pixel size |  | 
Yes& |  | 
352 Cells (full memory) |  | 
Mode I | Mode II | Mode III
4.5 MHz |  | < 4.5 MHz
12 keV* | 20 keV | 12 keV*
> 104 @12.4 keV (maximum achievable) |  | 
Full Frame |  | 
Yes 0.35 ph RMS Noise | Yes 0.21 ph RMS Noise | Yes$ < 0.21 ph RMS Noise
1/10 of the pixel size |  | 
Yes& |  | 
512 Cells (full memory) |  | 
LPD
#Same performance is expected for 7 – 12 keV
*Same performance is expected for 7 – 9 keV
$Feasibility to be confirmed
&Remove empty frames for different bunch patterns
Notes: LPD
Operation mode 1
E=12 keV (the same performance is expected at 7-9 keV)
RMS noise = 1200 e- (0.35 ph@12 keV)
Single photon sensitivity at 7-9 keV is not required
Max. nr of ph/pulse/pixel before saturation > 104 photons
Full frame readout
4.5MHz repetition rate (full memory, i.e. 512 memory cells)
Veto/data reduction capability to adjust to the pulse rate
Alignment precision – according to WP-75 definition (see “Detector Alignment at European XFEL: Concepts & Responsibilities” on Alfresco)
Operation mode 2:
E= 20 keV
0.21 ph noise @ 20 keV
Rest stays the same as in the mode 1
Operation mode 3:
E= 12 keV (the same performance expected at 7-9 keV)
Longer readout circle (no 4.5MHz speed) mainly for longer reset (noise reduction) if the feasibility study will show improvement in the detector performance (i.e. lower noise)
Rest stays the same as in the mode 1

## Slide 19: Detector Operation Modes – First Day
Non Linear System Response
DSSC
 | Mode I | Mode II | Mode III
Read Out Frequency | 4.5 MHz |  | 
Photon Energy | 1 keV | 1.5 keV | 0.7 keV
Max # of ph/pulse/pixel | 3 x 103 @1keV (maximum achievable) |  | 
Read Out Geometry | Full Frame |  | 
Single Photon Sensitivity | No |  | 
Alignment Precision | 1/10 of the pixel size (quadrant level) |  | 
Veto Capability | Yes& |  | 
Memory | 800 Cells (full memory) |  | 
AGIPD
Mode I | Mode II | Mode III
4.5 MHz |  | 
8.3 keV# | 15 keV | 7 keV
104 @12.4 keV (maximum achievable) |  | 
Full Frame |  | 
Yes 0.15 ph RMS Noise |  | No
1/10 of the pixel size |  | 
Yes& |  | 
352 Cells (full memory) |  | 
Mode I | Mode II | Mode III
10 Hz |  | 
1 keV | 1.5 keV | 0.7 keV
3.6 x 102 @1 keV (maximum achievable) |  | 
Full Frame |  | 
Yes |  | 
1/10 of the pixel size |  | 
Not available |  | 
Not available |  | 
FastCCD
#Same performance is expected for 7 – 12 keV
&Remove empty frames for different bunch patterns
Notes: LPD
Operation mode 1
E=12 keV (the same performance is expected at 7-9 keV)
RMS noise = 1200 e- (0.35 ph@12 keV)
Single photon sensitivity at 7-9 keV is not required
Max. nr of ph/pulse/pixel before saturation > 104 photons
Full frame readout
4.5MHz repetition rate (full memory, i.e. 512 memory cells)
Veto/data reduction capability to adjust to the pulse rate
Alignment precision – according to WP-75 definition (see “Detector Alignment at European XFEL: Concepts & Responsibilities” on Alfresco)
Operation mode 2:
E= 20 keV
0.21 ph noise @ 20 keV
Rest stays the same as in the mode 1
Operation mode 3:
E= 12 keV (the same performance expected at 7-9 keV)
Longer readout circle (no 4.5MHz speed) mainly for longer reset (noise reduction) if the feasibility study will show improvement in the detector performance (i.e. lower noise)
Rest stays the same as in the mode 1

## Slide 20: Calibration Plan & Schedule – Example DSSC
+1 year
+1 year
+1 year
+1 year
+1 year
+1 year
XFEL
Munich
XFEL
Munich

## Slide 21: Steps Towards First Day of Operation
Before delivery of the detector to XFEL
  Early performance testing and commissioning of prototype systems (single modules/quadrants) together with developer groups with in-house and external sources
  Control and DAQ integration as early as possible
  Participate in and perform performance optimization and calibration together with development groups
XFEL
AGIPD
DSSC
LPD
Delivery of the detector to XFEL
  Verification of detector performance, control system and DAQ/data management system “end-to-end” test in the XFEL detector laboratories at HERA South without XFEL beam
  Commissioning and initial calibration of the detector at HERA South before integration of the system into the beam line experiment
Detector Integration into and Commissioning at Beam Line
  DAQ and control system installation will take place before the detector will be installed
  Installation of the detector system and functional tests (vacuum, mechanics, control system, DAQ system, data management system) → repeat end-to-end test at beam line
  Commissioning and alignment with XFEL beam in-situ
    Detector ready for first day user operation

## Slide 22: Detector Tests and Commissioning – HERA South
End-To-End Test
Goal of the Test
Demonstrate the functionality of the detector system including all subcomponents and identify potential problems before system acceptance and integration at beam line.
What will be Tested?
Vacuum, cooling and movement system
Detector control system
C&C and veto system
DAQ and data management system including calibration data base
Data processing pipeline and calibration pipeline
Detector operation modes I-III for first day operation (priority)
Data and calibration data quality checks
Where and How will be Tested?
HERA South: BigAmber or PulXar depending on tests to be performed
Data taking with and without X-rays
Sequentially all four quadrants of a detector at full speed 4.5 MHz
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 23: Steps Towards First Day of Operation
Before delivery of the detector to XFEL
  Early performance testing and commissioning of prototype systems (single modules/quadrants) together with developer groups with in-house and external sources
  Control and DAQ integration as early as possible
  Participate in and perform performance optimization and calibration together with development groups
XFEL
AGIPD
DSSC
LPD
Delivery of the detector to XFEL
  Verification of detector performance, control system and DAQ/data management system “end-to-end” test in the XFEL detector laboratories at HERA South without XFEL beam
  Commissioning and initial calibration of the detector at HERA South before integration of the system into the beam line experiment
Detector Integration into and Commissioning at Beam Line
  DAQ and control system installation will take place before the detector will be installed
  Installation of the detector system and functional tests (vacuum, mechanics, control system, DAQ system, data management system) → repeat end-to-end test at beam line
  Commissioning and alignment with XFEL beam in-situ
    Detector ready for first day user operation

## Slide 24: Detector Tests and Commissioning – at Experiment (First Steps)
Without XFEL Beam
Requirements
Same conditions as for End-To-End test this time at experiment
Calibration source installed and operational
  Detector installed and safety interlocks fully operational
  No direct and not attenuated beam can hit the detector
  Tests to be done
  Detector performance verification measurements of Operating Mode I-III with and without calibration source under nominal operation conditions (pressure, temperature, sampling speed)
  Analysis of performance data and update of CalDB
With XFEL Beam
Requirements
Same conditions as for End-To-End test this time at experiment
Calibration source installed and operational
  Detector installed and safety interlocks fully operational
  No direct and not attenuated beam can hit the detector
  Tests to be done
  Alignment measurement with diffraction sample
  Detector performance verification measure-ments of Operating Mode I-III with XFEL beam
  Analysis of performance data and update of CalDB
  Final operation tests with up-to-date CalDB and alignment data
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 25: Commissioning – Base Line Operating Modes
Required Pulse Patterns
Time
100 ms (10 Hz)
600 μs
Pulse Train
Pulse Train
2700 Pulses
Time
100 ms (10 Hz)
220 ns or n x 220 ns
Pulse Train
Pulse Train
Time
100 ms (10 Hz)
220 ns
Pulse Train
Pulse Train
Time
100 ms (10 Hz)
Pulse Train
Pulse Train
Pulse Train
Nominal Pulse Pattern
Single Pulse per Train
Two or More Pulses per Train
One Pulse Every Second Train
(10 Hz Detectors)
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 26: Commissioning – Base Line Operating Modes
Time
100 ms (10 Hz)
600 μs
Pulse Train
Pulse Train
2700 Pulses
Nominal Pulse Pattern
Further Requirements
  Most important photon energies 0.7, 1, 1.5,
  7, 8.3, 12, 20 keV more are welcome
  Beam monitor accuracy for I0 measurement
  Beam intensity up to 105 ph/pixel/pulse
  Beam profile characterized
  Beam attenuators for intensity scans
  (gain and dynamic range validation)
  Unfocussed beam (d ≈ few centimeters or more) and focused beam (d ≈ 20 μm)
  Fluorescent target for e.g. flat field measurements and 55Fe source installed in experiment chamber
Stable Experiment Infrastructure
Vacuum, cooling and movement system
Detector control, C&C and veto system
DAQ/data management system including calibration data base
Data processing pipeline and calibration pipeline
Reference sample for alignment measurements
  Assumption
  Infrastructure is operational and tested up to the level that is possible without a detector.
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 27: Dynamic Range Scan and Gain Validation
  Implementations
  Multiple
  parallel amplification
  stages
  Gain switching or charge resetting mechanism
  Non linear component on front-end or sensor level (sensor, preamplifier)
Number of Incident Photons
Digital Output
Number of Incident Photons
Digital Output
Number of Incident Photons
Digital Output
Low
Medium
High
Mult. Linear Amplifiers + ADC
Gain Switching Amplifier + ADC
None Linear System + ADC
Image: AGIPD Collaboration
Image:RAL
Image: DSSC Collaboration

## Slide 28: Dynamic Range Scan and Gain Validation
Number of Incident Photons
Digital Output
Number of Incident Photons
Digital Output
Number of Incident Photons
Digital Output
Low
Medium
High
Mult. Linear Amplifiers + ADC
Gain Switching Amplifier + ADC
None Linear System + ADC
Gain | Source | Detector
High | PulXar/Panter | AGIPD, DSSC, LPD prototypes, FastCCD, SDD
 | Big Amber | AGIPD, DSSC, LPD prototypes and1 Mpixel, Gotthard V1/2
 | Little Amber | Gotthard V1 and II, LPD 2-tile system
 | Phoebe | DSSC prototype and 1 Mpix, FastCDD, pnCDD
Medium to Low | PulXar/Panter | DSSC prototype and 1 Mpix
 | LED Array | DSSC prototype and 1 Mpix LPD
 | Internal Calibration Circuit | AGIPD/LPD
 | Synchrotrons | AGIPD/DSSC/LPD prototypes
 | LCLS | AGIPD/DSSC/LPD prototypes
All | XFEL | Validation for all detectors

## Slide 29: Dynamic Range Scan and Gain Validation
Number of Incident Photons
Digital Output
Number of Incident Photons
Digital Output
Number of Incident Photons
Digital Output
Low
Medium
High
Mult. Linear Amplifiers + ADC
Gain Switching Amplifier + ADC
None Linear System + ADC
  Beam Time Estimate for Gain Scan
  Required data points per gain and per detector
  ≈10
  x3 gain stages
  x3 photon energies
  approx. 5h x 3 x 3 = 45 h ≈ 2 days
  Breaks possible, instrument team available
  Additional planned tests and measurements:
  EM compatibility and noise tests
  Systematic background studies (beam jet, stray light etc.)
  and more depending on requirements …
  These studies are presently only partially included in the installation schedule!

## Slide 30: Example: AGIPD 1 Mpx Integration at SPB
Status of Planning
Template resource loaded detector integration schedule exists for vacuum and non-vacuum detectors.
Detailed schedule and approach has been discussed with FXE, SPB and MID.
Milestones have partially been synchronized with scientific groups up to the level presently known.
Responsibilities are agreed with FXE, SPB, MID, CAS, ITDM and AE.
Definition of planning interface to PSPO and TC.

## Slide 31: Example: AGIPD 1 Mpx Integration at SPB
Open Issues
Agreement on responsibilities with CIE.
Establish planning interface to PSPO and TC.
Integration of the detector installation schedule into the XFEL “master plan” and resolve collisions.
Fix dates for incoming milestones with instrument groups, AE, ITDM and CAS.
Prepare planning for SQS, SCS, and HED and synchronize milestones.
So far no contingency is foreseen to be consistent with the global XFEL “master plan”.

## Slide 32: Example: AGIPD 1 Mpx Integration at SPB
Total Duration 8 Weeks
Commissioning with Beam
Commissioning without Beam

## Slide 33: AGIPD 1 Mpx Integration at SPB – Tasks

## Slide 34: AGIPD 1 Mpx Integration at SPB – Tasks

## Slide 35: Thank you for your Attention

## Slide 36: Bunch Trains
60 to 2700 bunches per train at a 10Hz repetition rate
78 to 3510 W beam power at 130MeV and 1nC (10-472kW at 17.5GeV)
Expected dark current of up to about 30 W at 130MeV and 600 μs
Orbit/energy stability along train (intra-train feedbacks)
Bunches from each train can be used for destructive semi-parasitic diagnostics (emittance, …)
B. Beutner & W. Decking XFEL Commissioning Meeting, 25.02.2015

## Slide 37: Injector Commissioning: Goal(s)
Achieve all XFEL parameters
Reduced/relaxed parameter set:
  beam charge around 0.5 nC good for diagnostics and LLRF
  Linac operation with several tens of bunches
    single bunch operation is useless for LLRF diagnostics and studies
    many bunch operation is more demanding for machine protection
Quantity | TDR goal | initial goal
macro pulse repetition rate | 10 Hz | 10 Hz
RF pulse length (useable flat top) | 600 ms | 600 ms
bunch repetition frequency within pulse | 4.5 MHz | 100 kHz
bunch charge | 0.02 – 1 nC | 0.5 nC
Slice emittance | 0.4 - 1.0 mm mrad | 1.0 mm mrad
Gun gradient | 60 MV/m | 50 MV/m
B. Beutner & W. Decking XFEL Commissioning Meeting, 25.02.2015

## Slide 38: Initial e- Beam Parameters – Bunch Repetition
Linac will be operated with several tens of bunches (not single bunch) as fast as possible
RF pulse length can be varied (gun performance and dark current losses might be limiting factors)
Initially possible repetition rates (injector laser)
Many bunches make European XFEL unique
Should be established as soon as possible
n | bunch repetition frequency [MHz] (=1.3GHz/(n*144) with n>=2 and even) | bunch spacing [ns] | max # of bunches in 600 μs train
2 | 4.514 | 222 | 2700
3 | 3.009 | 333 | 1800
8 | 1.128 | 886 | 670
90 | 0.1003 | 10000 | 60
- | - | - | 1
B. Beutner & W. Decking XFEL Commissioning Meeting, 25.02.2015

## Slide 39: Overview of e--Beam & FEL Operation Modes
Electron energy
8.5, 12, 14, 17.5 GeV
±1.5 % fast scanning
Bunch charge/compression -> X-ray pulse length
20-1000pC -> 2-100 fs
Long. Bunch shape
Bunch repetition –
X-ray delivery pattern
From pulse-on-demand to 4.5 MHz
Properties on demand – within bunch train
Special
Seeding
Tapering
Variable Polarization
Mulit-color
Instrument can determine
  Exact photon energy
  Focus size
  Bandwidth
Simultaneous operation
of 3 (5) experiments