## Slide 1
TCT measurements with CHESS-1 chip

## Slide 2
Periphery + signal to HV and readout
Only signal to readout:
(all pixels to HV):
Substrate
Periphery
Signal
Beam direction
100 um
45 um
no cross talk, no charge
from the low edge
Reminder: Top TCT on PPA9
(from 1st of September)
only n-well can be seen, other
are covered by metal
100 μm
45 μm

## Slide 3
Top -TCT with passive pixel array PPA10 on Chess 1
 no contact rings around pixels => not covered with metal
measured with chip irradiated to 2e15 with neutrons
Chess1
PPA10
substrate
periphery
signal

## Slide 4
Pixel length
A
Pixel width
B
NWELL length
NWELL width
nwell dimension ~ 51 um x 19 um
structure PPA10 has A = B = 0 (no substrate contact between n-wells).
substrate contact on the edge of the structure.

## Slide 5
signal to HV and readout via Bias-T
perimeter to HV
(and readout)
to ground
HV
scope
GND
substrate
connections
narrow (FWHM ~ 7 μm) laser beam
 measure with IR ( 1 mm absorption length) and red light (3 μm absorption length)
Laser beam
in all plots bias voltage was 120 V

## Slide 6
Red light (λ = 640 nm),3 μm absorption length
all pixels connected to readout
Charge
n-well: absorption in the well (less light into E field under the junction)
Bond pads and shade of bond wires
structure
border
not processed
area
Photo of the structure
Bond-wires

## Slide 7
50 μm
(n-well)
Pixel length 100 μm
Red light (λ = 640 nm),3 μm absorption length
all pixels connected to readout,
Zoom on one pixel:
20 μm

## Slide 8
IR light – 1 mm absorption length, simulates MIP
all pixels read out
 no large variations of collected charge across the surface

## Slide 9
only signal pixel read out
 in this structure 2 n-wells from 1 pixel
200 μm (two n-wells )
charge only from signal pixel
 no cross talk
Red laser
IR laser
red laser
reflection from bond wire

## Slide 10
IR lase, velocity (tells something about average electric field)
indicates bands of lower field between pixels
 gaps only in y direction  maybe becasue of larger distance between n-wells
in this direction?

## Slide 11
Summary
TCT measurement with narrow laser beam directed to the top of PPA10 structure
(irradiated with 2e15 neutrons, bias 120 V)
structure PPA10 not covered by metal
can clearly see n-wells
no large variation of charge collection across the surface with IR light
lower field regions between n-wells
would be interesting to measure this before irradiation