## Slide 1: ToF PET Peter Fischer, Heidelberg University 9th Terascale Detector Workshop Freiburg, 5.4.2016

## Slide 2: Overview
Introduction to PET and ToF-PET
Limits in Time Resolution with Scintillator Readout
It's all about PDE
Module Design & Results

## Slide 3: PET Scanners
e+ emitter is injected into body (18F, T1/2 ~ 2h)
e+ + e- → g + g
g (511 keV) pairs are detected in modules
Coincident (ns) hits → Pair → Line-of-Response 'LOR'
Backprojection gives 3D image
Modules often arranged in a ring
Note: 'pointing' geometry improves resolution only in centre of FOV
g
g
Field
of View

## Slide 4: Benefit by Time-of-Flight (ToF)
no TOF
t1 = 1.3 ns
t2 = 1.7 ns
with TOF
Dx = c/2 Dt
CRT = 0.4 ns

## Slide 5: Briefly: Definition of Time Resolution
In order to translate time resolution Dt directly to spatial resolution Dx:
CRT (Coincidence Resolving Time) = CTR (Coincidence Time Resolution) = FWHM of two channels in coincidence
CRT = FWHM2 = √2 FWHM1 = √2 × 2.35 s = 3.33 s
Remember:
  CRT = 100 ps ⟺ s1chan = 33 ps ⟺ FWHMX = 15 mm
For comparison: TDC with 100 ps time bins
  s = 100 ps / √12 ≈ 29 ps
t
1
½
FWHM

## Slide 6: ToF Benefit: Simulations @ Uni Ghent
PET scanner with d = 65 mm
GATE5 – Geant4 Monte-Carlo
Chart: NPW-SNR
| r = 50 mm | r = 100 mm | r = 150 mm
2 mm | 2.55 | 2.8 | 3.7
2.6 mm | 2.6 | 2.85 | 3.4
4 mm | 2 | 2 | 2.5
Chart: NPW-SNR
| r = 50 mm | r = 100 mm | r = 150 mm
200 ps | 3.95 | 4.55 | 4.1
400 ps | 2.95 | 3.3 | 3.7
600 ps | 2.5 | 2.8 | 3.45
no TOF | 1.65 | 1.8 | 2.3
H.Tohen et al, Phys.Med.Biol. 58 (2013) 6459-6479
Pixel size reduction gives benefits only down to about 2.6 mm
ToF has large benefit
NPW-SNR: non-prewhitening signal-to-noise ratio
Effect of crystal size
@ CRT = 600 ps
Effect of ToF
@ pixel = 2.6mm

## Slide 7: The Two Worlds of PET
PET Ring / Scanner
Clinical PET
(Whole Body PET)
- For humans
- large diameter Field Of View (>60 cm)
- spatial resolution: few mm
- time resolution CRT≪ 400 ps for ToF
- high sensitivity (low dose) → large area
Preclinical PET
(Animal PET)
- For mice, rats, rabbits (& human brain)
- Small diameter FOV (4-15 cm)
- spatial resolution: ≪ 1 mm
- time resolution only for coinc. (few ns)
- Depth-of-Interaction desirable to fight parallax effect

## Slide 8: Commercial PET Scanners
Nuklearmedizin 2012; 51 doi:10.3413/Nukmed-0455-11-12
Mediso Product Catalog
Mediso Product Catalog
Clinical
Preclinical
(Clinical)

## Slide 9: Readout Principles
Nearly all systems based on readout of (heavy) crystal scintillators:
Photo Detectors can be
  Photomultipliers + external electronics (past commercial systems)
  APDs or Analogue SiPMs + ASICs
  Digital SiPMs = 'SPAD' in CMOS ASICs (not considered here)
1:1 coupling
fast
Block detector
Light spreader
Spatial Resolving
Sensor
high position resolution / less channels
Tricks
depth of interaction (DOI)

## Slide 10: (Our) PET Modules
12 × 12 LYSO crystals in 2.5 mm pitch (10 mm height)
Full readout in 4 ASICs
Water cooled substrate

## Slide 11: Factors Affecting Time Resolution

## Slide 12: Shape of Light Pulse
Not nice at all! It’s too slow for known materials !
But why does decay time affect time resolution? - # of photons is fixed - peak amplitude decreases for slower pulses
 | NaI(Tl) | BGO | LSO | LYSO | LaBr3:Ce
Att. length @511keV [mm] | 29 | 10.4 | 11.4 | 11.4 | 21.3
Photofraction % | 17-20 | 40 | 32 | 30-35 | 20
Light Yield [photons / MeV] | 41000 | 9000 | 26000 | ≈ 25000 | 70000
Decay time [ns] | 230 | 60/300 (1/10) | 40 ns | 40-44 ns | 15 ns
Refractive Index | 1.77 | 2.15 | 1.82 | 1.82 | 2.25-2.4
Center Wavelength | 410 | 480 | 420 | 420 | 358
Hygroscopic? | yes | no | no | no | yes

## Slide 13: Photons Arrive One-by-One
When zooming into the first few 100 ps, photons arrive 'one after the other'
  4 Monte Carlo Events, 511keV in LYSO, tdecay = 42 ns, 40% overall PDE
# detected
photons
ns
s = 33ps
(⟺ CRT = 100ps)
Must maximize number
of (detected) photons!
Can improve resolution if
arrival times of multiple
photons are known
(see S. Brunner's talk)

## Slide 14: What is the Best Threshold ?
Intuitive Answer: The lowest possible, i.e. trigger on ‘first photon’
But: The first photon has a statistical jitter. Higher Thresholds are BETTER!
  Also, low thresholds are sensitive to dark counts -> lots of fake triggers, deadtime
S. Seifert, H.T. van Dam and D. R. Schaart: The Lower Bound on the Timing Resolution of Scintillation Detectors, Phys. Med. Biol 57, 1797-1814, 2012

## Slide 15: Photon Losses
A lot of light is trapped in crystal by total reflection
  nLYSO = 1.82 , nSi = 3.5
  Long / old discussion: specular or diffuse reflection ?
Photons are not absorbed in SiPM (geometric fill factor)
FBK, cells 67 x 67 um2

## Slide 16: More Photon Losses: Spectral Matching
Spectral response of SiPM must match emission spectrum of crystal
Need UV-sensitivity!
NUV-HD
RGB-HD
CsI(Tl)
LSO,
NaI
LaBr,
CeBr
25mm cell pitch
Over-voltage = 9 V
FBK devices:
NEW
generation
'OLD'
generation

## Slide 17: Long Crystal Effect
Consider a ‘long’ crystal (today: 4 x 4 x 20 mm3)
Absorption at top:
Photon travels long (60 ps for 10mm) (must compare to speed of gamma)
few photons arrive
Absorption at bottom:
short path
many photons arrive
Measurement:
Crystal 2 x 2 x 18 mm
rough sides, gap + white epoxy
polished end, PTFE reflector tape
Amplitude variation
leads to time walk

## Slide 18: Summary: Optimizing for Time Resolution
Maximize Number of Photons (in first 100 ps)
  Fast scintillator (LaBr: hygroscopic, bad abs.)
  Short Crystal (bad efficiency)
  Good optical coupling
Use a good SiPM
  High PDE of SiPM (fill factor, wavelength matching)
  Use a SiPM with large cells (fill factor) (bad energy resolution)
  Optimize SiPM for fast signal part (not discussed here)
  Optimize Dark Count Rate (low threshold, high overvoltage)
Measure @ cold (reduce dark count rate)
Optimize Threshold
It's all about the
overall PDE!

## Slide 19: It's all about PDE (1)
Very nice, recent results from FBK (from FAST workshop in March 16)
  Thanks to Alessandro Ferri & Claudio Piemonte for their ok to show a few results!
Comparison of SiPMs with different small cell sizes and (still high) fill factor:
'NUV' technology sensitive at short wavelengths:
  matching LYSO emission

## Slide 20: It's all about PDE (2)
CRT measurements with
  Small LYSO crystals of 3×3×5 mm3, t~40ns
  No SiPM 'overglass' protection
  Optical glue: melt-mount (n=1.54)
  5-10 layers of Teflon
  4×4 mm2 SiPMs
  Discrete amplifier or NINO (similar results)
  Readout with sampling Scope
  Energy cut at 511 keV ±FWTM
  Optimal Threshold for best timing:
  Conversion to Photoelectrons via 1/2/3… PE amplitude spectrum
Parameter:
Overvoltage

## Slide 21: It's all about PDE (3)
CRT vs. overvoltage for the 4 designs: Down to 100 ps!
How to explain
  Overvoltage dependence
  Differences between designs
FBK

## Slide 22: It's all about PDE (showdown)
Plot CRT vs. PDE (instead of overvoltage)
  PDE is measured vs. l and calculated for LYSO emission spectrum
FBK
FBK
Limit: perfect device
with PDE = 100%:
CRT would be 60ps

## Slide 23: For comparison: 'very small' crystals @ CERN
S. Gundacker et.al., NIMA, dx.doi.org/10.1016/j.nima.2013.11.025 (P. Lecoq group)
SiPM: MPPC (Hamamatsu, 3×3 mm2)
Crystal: LSO:Ce, 0.4% Ca (‘best’), 2×2×L mm3, wrapped in Teflon, optical grease
Readout: NINO ASIC
Strong effect of
crystal length

## Slide 24: Full Module Performance

## Slide 25: HyperImage vs. Sublima Detector Module
144 Channels
SiPM pitch 2.5×2.5mm2
Internal liquid cooling
64 Channels
SiPM pitch 4×4mm2
Hyperimage Project
Sublima Project

## Slide 26: 2nd Generation Module
Single ended FE → no more coupling components needed
ONE substrate only (low temperature co-fired ceramics, LTCC)
  SiPMs on top side
  PETA5 on bottom side
Liquid Cooling inside the substrate (MSE, Germany)
PETA Chips
Flip chip mounting
MR compatible components and layout
SiPM side

## Slide 27: Sensor Geometry
Active area per SiPM 2.25 × 2.25 mm² ('OLD' design, cell size 25×25 µm²)
Rectangular dies with 2 × 6 SiPMs
Wire-bonding pads at the external edge → no gaps in the active region
5mm
15mm
perfectly regular 2.5mm pitch all over the array
Device from FBK

## Slide 28: Bumped Chips
Wire bonds take much space → use solder balls in 270 µm pitch
UMC 180nm technology (6 layers)
5 x 5 mm² size
5 × 5 mm2

## Slide 29: PETA4-6 Channel
I→I
Ubias
DAC
ADC
Disc
∫i(t)dt
Disc
Thres
DAC
MUX
Channel
Logic
TDC
∫u(t)dt
Single Ended FE
Differential FE
Neighbor
triggers
I→U
ADC
Filter
IN+
IN-
INSE
Hit Logic
control
readout
control
Two FE versions per channel:

## Slide 30: Module Time Resolution
Two 144 channel modules facing each other
1:1 coupling, 10 mm long LYSO, room temperature
22Na

## Slide 31: 1 channel vs. 1 channel
  One channel per side. Move source along the line of response
Δx = 0.5 × ΔT× c
15mm 15mm
100ps 100ps

## Slide 32: Many Channels ('All vs. All')
'Problem': Our point source only creates coincidences in few crystal combinations
Must therefore use many positions and overly results.
(This explains badly filled 'channel - channel' bins)
A1
A8
B1
B8
A1
A8
B1
B8

## Slide 33: Timing resolution: '18 vs. 18'
OV = 8V (only), 22°C
missing source positions
broken channel
µ ~ 232 ps

## Slide 34: Comparison to FBK results ('gold standard')
SiPM bias voltage scan
Results very close to the limit for 'OLD' RGB-HD technology
Overvoltage can be increases much higher for single crystals
Limitation in array operation also observed in FBK discrete setup!
FBK 64 Ch 10°C
FBK 1 Ch 10°C
PETA 36Ch 22°C
PETA 36Ch 34°C
SiPM 4x4 mm2
LYSO 3x3x5 mm3
Measurements with discrete electronics
SiPM 2.5x2.5 mm2
LYSO 2.5x2.5x10 mm3
200 ps
300 ps
180 ps
PETA @ RT

## Slide 35: Summary: Module Performance
Full modules with 'OLD' SiPMs from FBK have an average CRT of 230 ps
  (single side gamma time resolution s = 70ps)
When comparing to single channel measurements, must consider
  We use with off-the-shelf LYSO
  SiPM (RGB-HD) is not 'latest' SiPM generation
  Many SiPMs operate in parallel
  Crystals are rather small (2.25 mm) (10 mm long)
  Measurement done at room temperature
  Crystal isolation (ESR / Al / ESR) is (probably) worse than Teflon
Spatial resolutions in preclinical arrangements are excellent (<0.8mm)

## Slide 36: Summary
ToF improves image quality (or: lowers dose, or: lowers scan time) in PET
Commercial devices have 400-600 ps CRT
CRT close to 200 ps @ RT has been demonstrated (by several groups)
  Multi channel R/O ASIC can cope with this resolution
Progress in SiPMs (fill factor, spectral sensitivity, dark noise) can bring us close to 100 ps
It's all about PDE
  Optimize SiPMs
  Maximize light (geometry, coupling, wrapping,..)
Need faster crystals!

## Slide 37: Thank you for your attention!

## Slide 38: Example of Improvement by ToF
Slide from A. Ganin, GE Company, CRT ≈ 400 ps, Hot Rods Phantom
Contrast enhancement / resolution increase (@ given dose)
C. Kim et al., ToF PET Detector development with Silicon Photomultiplier, 2012 IEEE NSS and MIC Record

## Slide 39: Energy Resolution (@ low overvoltage)
Difficult to determine exactly because FE saturates at high amplitudes
1.275 MeV peak cannot be used well for calibration and nonlinearity correct.
~12 %

## Slide 40: Speed of Rising Edge
Has been studied (TU Delft) directly with triggered X-ray pulses
Rise time ≈ 100 ps for LSO:Ce wrapped in black tape (only ‘direct’ photons)
But: White Teflon wrapping (‘to get all the light’ to improve energy resolution) worsens rise time!
Depends very much on Crystal Geometry

## Slide 41: SiPMs from FBK + Our Chip (PETA3)
crystal 2
crystal 1
Positron
Source
ASIC
Discrete setup with
‘best’ signal processing

## Slide 42: Requirements for Electronics
Low noise obvious
Low system threshold this is the real point…
High bandwidth obvious, but be aware of noise!
High pass filter to eliminate baseline fluctuations
Low input impedance (some 10 W) to get a large ‘fast’ SiPM component
Adjustable DC input potential to adjust overvoltage of SiPM
TDC bins ≈ 50ps quadratically < target CRT
Energy resolution 7-8 Bit to cut on 511 keV in 1:1 coupling. More bits for position reconstruction with light sharing
Fast Abort for false triggers on DCR to avoid large dead time
Hit selection from amplitude (on too much or little light, in 1:1)
Neighbor trigger in light sharing application
Data readout rate ≈ 100 khits / channel
Acceptable power dissipation (continuous operation)
Low Emission in PET / MRI

## Slide 43: Summary: PETA chips
PETA is a complete 38-Channel Timing Solution (incl. TDC w. 50 ps bins)
PETA allows compact module designs (flip chip, few external parts)
Choice between Differential FE (both polarities, MRT immune) and Single Ended FE (low Zin, DC bias adjustment, no external coupling parts)
Channel dead time < 500ns
Readout rates >200 kHz per channel (in all channels)
Power consumption ~30mW / channel (😥)
Special features:
  Neighbour logic (for light sharing concepts)
  Fast abort on 'noise hits' (small signals, before digitization, ⪅100ns)
  Abort on low amplitudes (≪ 511 keV, after digit., ~300 ns)
  Power down input (@ over-temperature, external measurement)

## Slide 44: Single Ended Frontend
Discriminator
Front-End
Inject
DAC
Charge Integrator & ADC
ThN
Finite State Machine (FSM)
VCM
Input
IRamp
- HV
SiPM
Inject
x 4
ThP
gain
VREF
Initial Condition
TDC
hit
clear
ampl.
time
No more external
components needed
Zin~ 7Ω
Variable DC
bias @ input
(~0.8V, per ch.)
Same as diff. FE
ADC 'gain'
ADC 'resolution' vs. conversion time
A low power front-end architecture for SiPM readout with integrated ADC and multiplexed readout, I. Sacco et al., JINST, Volume 8, January 2013, C01023
(I. Sacco)

## Slide 45: Readout
Fully self triggered
2 serial outputs per side (2 'spare' if higher rate needed)
Size of one event ~32 Bit
Rates well above 100 kHz simultaneously in each channel are possible
Improved PETA6 architecture has very low dead time:
Theoretical losses for
front-end dead time of 300ns
Simulated losses when using
1 output per side, @ slow clock
Output link capacity reached

## Slide 46: FE Chips for SiPM Readout
Requirements vary depending on application
  speed, dynamic range, #channels, resolution,…
So far only few system chips with all features ‘on board’
Overview Table from
Wojciech Kucewicz, AGH
And more:
PETIROC
RatCAP TSMC 180 32
STiC TSMC180 64

## Slide 47: SiPM Model
SiPM = parallel connection of many small SPADs
  Size ≈ 50um  N = 6400 for an area of 4x4 mm2
SPADs are biased with individual resistors to a DC potential above the breakdown voltage (20-30V)
When a SPAD is fired by photon, its voltage drops by the Overvoltage
DV
To understand output signal, caps must be added
(N-NFire)CD
(N-NFire)CQ
CQ
CG
CD
RQ
RQ/(N-NFire)
Firing cells

## Slide 48: SiPM Signal and Input Impedance
Current pulse @ resistive input has 2 components
  Fast component by capacitive coupling via CQs
  Slow component from recharging of CDs via RQs
The fast component of the input current becomes larger when input impedance is small!
For timing we NEED this component  we need small Rin
  For very small Rin (<10W) we are limited by parasitic Rs
(N-NFire)CD
(N-NFire)CQ
CQ
CG
CD
RQ
RQ/(N-NFire)
Zin
vin
isig
NFire times
SiPM Model
Amplifier
many 10 pF
few 10 fF
some 10 fF
DV

## Slide 49: Comparing Results
Comparing results from different groups / setups is difficult:
  Scintillator material (selected crystals? Co-doping? LaBr?)
  Scintillator Geometry (small cubes? long needles as required in PET?)
  Wrapping? (Many layer teflon not feasible in arrays? Black wrapping for timing?)
  Operation Temperature (cold allows higher overvoltage)
  Overvoltage / Dark noise rate (high OV gives better CRT, but high Darknoise and is often not feasible in practice due to dead time)
  SiPM Size