## Slide 1: Status of the Detector Development Program M. Kuster for WP75
May 23, 2016 20th Meeting of the European XFEL Detector Advisory Committee European XFEL GmbH, Hamburg

## Slide 2: Changes in CAS and Impact on WP75
B. Heisen resigned from group leadership position of CAS and as Karabo chief architect.
As interim solution the Karabo architecture tasks have been transferred to S. Hauf (WP75 Software and Analysis)
Works part time for WP75 in order to keep his activities there alive and stays affiliated with WP75
Consequences for WP75 to be evaluated as soon as Karabo plans and priorities are defined
Agreed with XFEL management
Review of the scope and time line of the measure by the ma-nagement board in August 2016 and decision by the MB about the next steps

## Slide 3
Adaptive Gain Integration Pixel Detector – AGIPD
Achieved Progress (Management)
  Monthly project status reports (November 2015 – April 2016) have been delivered to XFEL
Contract about the 2nd 1MPix + spare parts has been signed
Cost neutral prolongation for 1st AGIPD detector finalized and signed
Project review with XFEL management April 26th
AGIPD Single Module in N2 Box at HERA South
Achieved Progress (Technology)
Basic interlock concept defined
Dry N2 box produced and cooling installed in HERA South (for operation of module at low temperatures)
Motion system tests are planned for June and preparation of control hardware is ongoing
C&C tests with single module in preparation
Veto test setup with LED prepared and first tests are in progress (feasibility demonstrated)

## Slide 4
DePFET Sensor with Signal Compression – DSSC
Achieved Progress (Management)
Project report 12(1) + updated project plan and DSSC Phase II D19: Draft calibration plan delivered to XFEL (work on update is in progress)
Project status reports have been delivered to XFEL
Project review with XFEL management April 26th
Achieved Progress (Technology)
Mechanics design of the vessel is about to be completed
The DSSC ladder vacuum test stand, FENICE, assembled, commissioned and operational
Vacuum qualification measurements of DSSC ladder components
and other general components routinely ongoing
DSSC 1 Mpx MiniSDD camera integration plan and schedule is
under preparation
Beam times at PETRA III and LABEC are under preparation
DSSC mechanics
FENICE and Ladder Module

## Slide 5
5
Large Pixel Detector – LPD
Achieved Progress (Management)
Project status reports have been delivered to XFEL
Project review with XFEL management April 19th
Achieved Progress (Technology)
Design of revised detector housing reviewed and approved – implementation started
Beam time at ESRF and APS analysis of data is finished
Demonstrated improved noise performance with new power card.
End-to-End test preparations are in full swing
Test of internal charge injection circuit for gain
calibration is in progress
Final Installation of LPD at FXE
Si Powder Diffraction Pattern in x10 and x1 gain

## Slide 6
FastCCD
Achieved Progress
Camera assembled and tested (works with hole)
Data from tests available and being analyzed at XFEL
Camera to arrived at XFEL last week, commissioning is in progress
Detector integration into SCS experiment
Interface for SCS produced at LBNL and integration into SCS is under preparation at XFEL
Additional cooling adapter for more flexibility being designed at XFEL
Vacuum chamber
Assembled Camera & Chip
Fe55 Image
Fe55 Spectrum

## Slide 7: Detector Laboratory Munich
Achieved Progress (Management)
Rental contract for offices and laboratory signed
Application for usage of radioactive sources finalized
Achieved Progress (Technology)
Offices and laboratory refurbished
Munich group moved to Siemens Campus
IT infrastructure is installed
Small DSSC prototype test set-up working
Office and Laboratory in Munich (Siemens Campus)

## Slide 8: Reminder Base Line Detector Operation Modes – First Day
DSSC
 | Mode I | Mode II | Mode III
Read Out Frequency | 4.5 MHz |  | 
Photon Energy | 1 keV | 1.5 keV | 0.7 keV
Max # of ph/pulse/pixel | 3 x 103 @1keV (maximum achievable) |  | 
Read Out Geometry | Full Frame |  | 
Single Photon Sensitivity | No |  | 
Alignment Precision | 1/10 of the pixel size (quadrant level) |  | 
Veto Capability | Yes |  | 
Memory | 800 Cells (full memory) |  | 
AGIPD
Mode I | Mode II | Mode III
4.5 MHz |  | 
8.3 keV# | 15 keV | 7 keV
104 @12.4 keV (maximum achievable) |  | 
Full Frame |  | 
Yes 0.15 ph RMS Noise |  | No
1/10 of the pixel size |  | 
Yes |  | 
352 Cells (full memory) |  | 
Mode I | Mode II | Mode III
4.5 MHz |  | < 4.5 MHz
12 keV* | 20 keV | 12 keV*
> 104 @12.4 keV (maximum achievable) |  | 
Full Frame |  | 
Yes 0.35 ph RMS Noise | Yes 0.21 ph RMS Noise | Yes$ < 0.21 ph RMS Noise
1/10 of the pixel size |  | 
Yes |  | 
512 Cells (full memory) |  | 
LPD
#Same performance is expected for 7 – 12 keV
*Same performance is expected for 7 – 9 keV
$Feasibility to be confirmed

## Slide 9: Reminder Base Line Detector Operation Modes – First Day
DSSC
 | Mode I | Mode II | Mode III
Read Out Frequency | 4.5 MHz |  | 
Photon Energy | 1 keV | 1.5 keV | 0.7 keV
Max # of ph/pulse/pixel | 3 x 103 @1keV (maximum achievable) |  | 
Read Out Geometry | Full Frame |  | 
Single Photon Sensitivity | No |  | 
Alignment Precision | 1/10 of the pixel size (quadrant level) |  | 
Veto Capability | Yes |  | 
Memory | 800 Cells (full memory) |  | 
AGIPD
Mode I | Mode II | Mode III
4.5 MHz |  | 
8.3 keV# | 15 keV | 7 keV
104 @12.4 keV (maximum achievable) |  | 
Full Frame |  | 
Yes 0.15 ph RMS Noise |  | No
1/10 of the pixel size |  | 
Yes |  | 
352 Cells (full memory) |  | 
Mode I | Mode II | Mode III
10 Hz |  | 
1 keV | 1.5 keV | 0.7 keV
2.5 x 102 @1 keV (maximum achievable) |  | 
Full Frame |  | 
Yes |  | 
1/10 of the pixel size |  | 
Not available |  | 
Not available |  | 
FastCCD
#Same performance is expected for 7 – 12 keV

## Slide 10: Reminder Commissioning – Pulse Patterns
Required Pulse Patterns
Time
100 ms (10 Hz)
598.4 μs
Pulse Train
Pulse Train
2700 Pulses
Time
100 ms (10 Hz)
221.5 ns or n x 221.5 ns
Pulse Train
Pulse Train
Time
100 ms (10 Hz)
221.5 ns
Pulse Train
Pulse Train
Time
100 ms (10 Hz)
Pulse Train
Pulse Train
Pulse Train
Nominal Pulse Pattern
Single Pulse per Train
Two or More Pulses per Train
One Pulse Every Second Train
(10 Hz Detectors)

## Slide 11: Commissioning – Start Up Operation
Start-Up 1.128 MHz Operation
Time
100 ms (10 Hz)
598.4 μs
Pulse Train
Pulse Train
2700 Pulses
Time
100 ms (10 Hz)
4x 221.5 ns = 886 ns spacing
Pulse Train
Pulse Train
Nominal Pulse Pattern
Multiple Pulses per Train 60-70
Minimize calibration effort  operate detectors at nominal speed of 4.5 MHz
Use veto capability to adapt to different pulse patterns (discard empty frames to reduce data volume)
Full veto functionality required for end-2-end test, commissioning phase and day-1 operation!

## Slide 12: Reminder – Detector Tests at HERA South
End-To-End Test
Goal of the Test
Demonstrate the functionality of the detector system including all subcomponents and identify potential problems before system acceptance and integration at beam line.
What will be Tested?
Vacuum, cooling and movement system
Detector control system
C&C and veto system
DAQ and data management system including calibration data base
Data processing pipeline and calibration pipeline
Detector operation modes I-III for first day operation (priority)
Data and calibration data quality checks
Where and How will be Tested?
HERA South: Big Amber
Data taking with and without X-rays
Sequentially all four quadrants of a detector at full speed 4.5 MHz
What is required?
Vacuum, cooling and movement system
Detector control, C&C and veto system
DAQ/data management system including calibration data base
Data processing pipeline and calibration pipeline

## Slide 13: End to End Test – Test Scenario
  Step 1: Transport of detector to HERA South
  Mechanically install detector, connect to power supplies and control system, functional tests of all components
  Step 2: Detector vacuum system tests
  Pump system to nominal operation pressure, cool detector, determine leak rate of vessel, RGA analysis
  Step 3: Test of safety system and procedures
  Detector shutdown procedures, SIB/interlock logic or other safety mechanisms implemented in the detector
  Step 4: Detector motion system tests
  Test of control system interplay with motion mechanics, limit and safety switch tests, maintenance positions, operation positions etc.
  Step 5: Timing and Clock & Control system tests
  Test of C&C system, synchronisation and veto mechanism
  Step 6: Performance verification measurements without X-rays
  Test of interplay with data bases, control/DAQ system, data taking macros, performance of baseline operation modes (dark, noise, bad/noisy/dead pixels, internal calibration circuit …), analysis of data …
  Step 7: Measurements with X-rays
  Test of interplay with data bases, control/DAQ system, analysis of data, performance measurements as far as possible with continuous X-ray source
AGIPD 1M Installed in Big Amber
AGIPD X-ray Window Flange

## Slide 14: End to End Test – Preparation in HERA South
  LPD/AGIPD DAQ and power cables: specified and procurement is in progress
  Support table for LPD/AGIPD 1Mpix in Big Amber: designed, components are ordered and assembly is ongoing
  C&C System: installed in Hera South and first tests with LPD were successful
  Macros: first LPD test are adapted and implemented
  Veto system: first tests were performed and test setup for LPD/AGIPD is in preparation
  AGIPD control system: power supplies integrated in Karabo and hardware tested
  AGIPD laboratory vacuum system: procurement of vacuum components on-going and control system requirement document finalized (different to beam line)
  AGIPD laboratory cooling system: design finalized and procurement is in progress
  AGIPD motion system: design of the final control system is in progress
AGIPD 1M Installed in Big Amber

## Slide 15: Detector Integration Milestones SASE I – III
AGIPD
Ready for First Experiments at SPB
SPB Instrument User Operation
AGIPD
End-To-End Test
HERA South
FXE Instrument User Operation
LPD
Ready for First Experiments at FXE
LPD
End-To-End Test
HERA South
Start AGIPD
Installation at SPB
Start LPD
Installation at FXE
LPD 1 Mpixel Detector at FXE
AGIPD 1 Mpixel Detector at SPB
AGIPD 1 Mpixel Detector at MID
MID Instrument User Operation
AGIPD
Ready for First Experiments at MID
2nd AGIPD
End-To-End Test
HERA South
Start AGIPD
Installation at MID
Delivery 1st AGIPD
Delivery LPD
Delivery 2nd AGIPD

## Slide 16: Detector Integration Milestones SASE I – III
AGIPD
Ready for First Experiments at SPB
AGIPD Instrument User Operation
AGIPD
End-To-End Test
HERA South
FXE Instrument User Operation
LPD
Ready for First Experiments at FXE
LPD
End-To-End Test
HERA South
Start AGIPD
Installation at SPB
Start LPD
Installation at FXE
LPD 1 Mpixel Detector at FXE
AGIPD 1 Mpixel Detector at SPB
AGIPD 1 Mpixel Detector at MID
MID Instrument User Operation
AGIPD
Ready for First Experiments at FXE
2nd AGIPD
End-To-End Test
HERA South
Start AGIPD
Installation at MID
1st AGIPD 1 Mpixel Detector
  31.08.2016 Start end-to-end test in HERA South
  +12 weeks Finish end-to-end test in HERA South and start installation at SPB
LPD 1 Mpixel Detector
  01.10.2016 Start end-to-end test in HERA South
  +12 weeks Finish end-to-end test in HERA South and start installation at FXE
2nd AGIPD 1 Mpixel Detector
  January 2017 start end-to-end test in HERA South (to be confirmed)
  +12 weeks Finish end-to-end test in HERA South and start installation at SPB
European XFEL urgently needs reliable detector delivery dates based on realistic estimates!!!
Delivery 1st AGIPD

## Slide 17: 21th XDAC – Organization and Dead Lines
Format of the Meeting
  More development projects
  More time for the commit-tee
Deadlines for the 21st Meeting of the DAC
  Submission of project reports
  November 11th 2016
  Submission of presentations
  December 7th 2016
  DAC Meeting
  December 14th till 16th 2016
  Avoid splitting of software/hardware sessions
  No parallel sessions
  Reduce session length

## Slide 18: Detector Group at the European XFEL
Thank you for your attention!