## Slide 1: Common approaches to data reduction in time-resolved XDS on liquids
Dmitry Khakhulin, FXE instrument (WP-81)

## Slide 2
XDS
XES
Diode
XFEL beam
Optical laser beam
XAS
Courtesy by European XFEL/Rey.Hori
A Suite of Simultaneous X-Ray Tools available:
- XAS
- Non resonant XES
- Resonant XES (RIXS)
- X-Ray Raman Scattering
- Wide(Small)-Angle X-ray Scattering
- …
Towards a High-Speed Molecular Camera for tracking chemical reaction dynamics at FXE
Struct. Dyn. 2, 054301 (2015)

## Slide 3: Liquid WAXS for probing the solvation
No monochromatization needed  pink beam compatible
Moderate focusing requirements  < pixel size (0.5 mm)
High repetition rate
High dynamics range (single photon105 photons/pixel)
High energy scattering - unique for XFEL
Reliability of the WAXS difference signal – optimization of the experiment
Raw data
Wide-Angle X-ray Scattering (WAXS) delivers global geometric structural dynamics of the solute and the surrounding solvent
Scattering from 10 um flat water jet,
1012 photons/pulse @ 15keV
Sample-to-detector distance120mm
Notes: Systems:
  1 Mpix with adjustable central hole (in construction)
  2-Tile prototype: 32 x 256 pixels.
  Super-module prototype: 256 x 256 pixels.
Technology: Hybrid pixel detector constructed from 4-side buttable sensor modules. 86% active area.
Pixel size: (500 mm)2, 500 mm thick Si.
Energy range: 5 – 20 keV.
Frame rate: 4.5 MHz, output rate: 5120 images/s.
512 images analogue on-chip storage.
Read noise: 1000 e- rms, referred to sensor input.
Dynamic range: 105 12 keV photons
Amplifier architecture: 3 parallel gain settings merged off chip within DAQ system, 0.03 lsb/ph to 30 lsb/ph.

## Slide 4: Solvation dynamics in Fe(II) complexes
Lawson Daku, L. M. J PCL, 2010, 1, 1830
[FeII(bpy)3]2+/ H2O
cage+solute
solvent
solute
DS (Q, t) = DSsolute(Q, t) + DScage(Q, t)
Differential diffuse scattering:
Difference signal DS (Q, t) = S(Q, t > 0) - Sref(Q, t < 0)

## Slide 5
Nax
Neq
LS->HS
http://henke.lbl.gov/
Scattering from 10 um flat water jet,
1012 photons/pulse @ 20keV
Sample-to-detector distance100mm
10um jet
HR WAXS with LPD
10 cm

## Slide 6
Scattering at 1012 photons/pulse incoming (20keV), detx=100mm,
Noise ~0.4 photons, space angle corrected I/cos3(2*theta)
for the pattern periphery (radius = 450 pixels)
Poisson’s distribution for counting:
for the pattern maximum (radius = 40 pixels)
Notes: High energy scattering will quickly become popular and benefit from the reprate

## Slide 7: Spin crossover in aqueous Fe(II) complexes
Nax
Neq
LS HS
10um jet
10 cm
Aqueous solution, 15 mM
10 um jet, 50% excitation fraction
1012 ph/pulse@20keV
1 pulse
100 pulses

## Slide 8: Preparing the acquisition: Figure Of Merit (FOM) scans
Acetonitrile
Water
The most critical step of the experiment -
optimization of experimental parameters
laser/x-ray overlapped in space at well positive delay
temporal overlap
a laser power scan
This signal is always there no matter if sample works or not,
it just has to absorb some laser light

## Slide 9
Possible workflow for optimization of exp. conditions
Raw images set
~ 1000 x 2D
~ 4 GB
Dark images set
~ 10 x 2D
~ 40 MB
Masked and Corrected (optionally) images
~ 1000 x 2D
~ 4 GB
~ 1000 x 1D
~ 4 GB
FOM calculation
and averaging
~ 1 x 1D
~ 10 kB
Basic normalization
(optional)
Fastest possible feedback needed,
Drop most of the expensive corrections,
i.e. no ADU->photons conversion,
only basic normalization and image calibration
A few images is usually enough, drop frames
Robustness and reliability is preferred over exact accuracy, as the signal is typically very strong
The most important first step, defines the success of entire experiment
Signal optimization

## Slide 10
Typical data processing workflow
via Tim van Driel (DTU)
Raw images set
~ 10000 x 2D
~ 40 GB
Dark images set
~ 250 x 2D
~ 1 GB
Masked and Corrected (2D ) images
~ 10000 x 2D
~ 40 GB
Azimuthal
Integration
+ 1D corrections
~ 10000 x 1D
~ 100 MB
Normalization
~ 10000 x 1D
~ 100 MB
Filtering and
outlier rejection
~ 8000 x 1D
~ 80 MB
Producing difference
(OFF subtraction)
~ 5000 x 1D
~ 50 MB
~ 100 x 1D
~ 1 MB
Averaging
TimingTool
Resorting (optional)
Analysis
Whenever possible keep the raw images:
non-assembled/corrected, in ADU
LPD
laser
X-ray
dark
dark

## Slide 11: Scattering corrections
Chart: theta(deg)
d × P
d
P (horiz.)
1D corrections
- Space angle (geometry)
- Sample absorption (if normal)
- “Phosphor” correction
2D corrections
- Flat-field (gain map), CM corrections
- Distortion (image in lab. frame)
- Dark subtraction
- Polarization correction
- Sample absorption (if not normal)
2q
2q
D
sample
Dp
1D corrections

## Slide 12: Fast Azimuthal Integration – PyFAI
Python Implementation
(4M image example, no readout overheads):
- Image slicing (loop+mean) 40 s
- Histogram (scatter operation - no parallelization)
cheap memory-wise 800 ms
- Look-up-table LUT (“scatter to gather” transformation)
memory costly 150 ms
on GPU 15 ms
- Compressed sparse row (CSR) representation of
LUT – optimal for GPU parallelization and memory
consumption <150 ms
on GPU 6 ms
Features
- Pixel splitting implemented - expensive but important for
large pixels like on LPD
- Supports masking, dark and polarization corrections
- Seemingly supports non-regular pixel grids of tiled detectors,
h5 in NeXus definition
- Optimized for GPU parallelization, good performance
- Python library (C-based ) of functions relevant for scattering
Developed by J. Kieffer (ESRF)
Goal – 200 Hz

## Slide 13
Data processing workflow for anisotropic scattering

## Slide 14
Raw images set
Data processing workflow for anisotropic scattering
via Tim van Driel (DTU)
~ 10000 x 2D
~ 40 GB
Dark images set
~ 10000 x 2D
~ 40 GB
Masked and Corrected (2D ) images
~ 10000 x 2D
~ 40 GB
Normalization
Filtering and
outlier rejection
OFF subtraction
Averaging
TimingTool
Resorting (optional)
Analysis
Azimuthal
segmentation
and integration
~ 10000 x 2D
~ 40 GB
~ 8000 x 2D
~ 32 GB
~ 5000 x 2D
~ 20 GB
~ 100 x 2D
~ 40 MB
~ 100 x 2 x 1D
~ 2 MB

## Slide 15
Summary
Fast user feedback is crucial (FOM scans)
The raw images (in ADU) should be kept and provided to users
The standard XDS data reduction loops (to 1D) should exist
and be available for users
Ideally, more advanced reduction tools like e.g. for anisotropic scattering or scattering during phase transitions
Thank you for attention!