## Slide 1: Accurate beam position calculation for large off-center positions
Benoît Roche  ESRF, Diagnotics
DEELS 2016
With the participation of K. Scheidt
Special thanks to Guenther Rehm for his great help!

## Slide 2
Introduction
Finding beam position with a numerical solver
  Method
  Numerical solvers available
Finding beam position with a lookup table
  Method
  Non-converging situations and solution
  Results
Future developments

## Slide 3: Introduction
The / formula suffers from increasing dislinearity for off-centered beam positions.
A more accurate position calculation is sometime required:
During the early phases of commissioning of our future ring,
For special studies of the lattice that implies large orbit deviations.
 the grey grid represents the real beam positions
 the red grid represents the corresponding calculated values with / formula
Example with one of the future BPM blocks of the ESRF-EBS storage ring:

## Slide 4: Introduction
The / formula suffers from increasing dislinearity for off-centered beam positions.
A more accurate position calculation is sometime required:
During the early phases of commissioning of our future ring,
For special studies of the lattice that implies large orbit deviations.
Horizontal and vertical cuts in the previous plot.
For very off-centered positions the error with / formula can be a factor of 2!
Example with one of the future BPM blocks of the ESRF-EBS storage ring:

## Slide 5
Introduction
Finding beam position with numerical solver
  Method
  Numerical solvers available
Finding beam position with a lookup table
  Method
  Non-converging situations and solution
  Results
Future developments

## Slide 6: Finding beam position with numerical solver
A numerical solver gives V = (VA, VB, VC, VD) knowing beam’s position (X, Y)
To do the reverse calculation, we can use a converging iterative process
Lets reduce the number of dimensions to see how it works:
Va
Vb
/ formula
numerical simulation
possible measurement

## Slide 7: Finding beam position with numerical solver
Using numerical simulation we can find beam’s position with an iterative process (Newton’s method):
Va
Vb
V
u
First guess: X1
X2 = X1 + V . u
u is a local change of V for X = 1
Measured beam position

## Slide 8: Finding beam position with numerical solver
Using numerical simulation we can find beam’s position with an iterative process (Newton’s method):
Va
Vb
V
u
First guess: X1
X2 = X1 + V . u
u is a local change of V for X = 1
Measured beam position

## Slide 9: Finding beam position with numerical solver
Using numerical simulation we can find beam’s position with an iterative process (Newton’s method):
Va
Vb
X1
X2
X3
X4
X5
The solution is considered as founded if
|Xn – Xn-1| < 1 nm

## Slide 10: Finding beam position with numerical solver
This method gives a value for incoherency if a perfect solution cannot be found:
Va
Vb
X1
X2
X3
X4
X5
position
incoherency

## Slide 11: Finding beam position with numerical solver
A 2D solver written with Matlab (G. Rehm)
References:
Sensitivity calculation of beam position monitor using boundary element method Tsumoru Shintake, Masaki Tejima, Hitoshi Ishii, Jun-ichi Kishiro DOI: 10.1016/0168-9002(87)90496-7
http://www.lnf.infn.it/acceleratori/dafne/NOTEDAFNE/CD/CD-10.pdf DANE TECHNICAL NOTE CD-10, ANALYSIS OF THE DANE BEAM POSITION MONITOR WITH A BOUNDARY ELEMENT METHOD A. Stella
Easy to implement
Quite fast
Cannot simulate complicated geometries (like stripline)
Not fast enough for real time processing of TBT data

## Slide 12: Finding beam position with numerical solver
A commercial solver like CST
Can simulate complicated geometries
Expensive
Long computation time

## Slide 13
Introduction
Finding beam position with a numerical solver
  Method
  Numerical solvers available
Finding beam position with a lookup table
  Method
  Non-converging situations and solution
  Results
Future developments

## Slide 14: Finding beam position with a lookup table
The method is the same, the only difference is that we compute a table which gives the simulated (VA, VB, VC, VD) for an array of beam’s position (X, Y). We use this table instead of using the solver during the iterative process.
Advantages:
You can afford to have long computation time because you just do the computation once,
It is easy to change the solver (for instance to increase its precision) by just updating the look-up table,
It is, in principle, the quickest method you can think of.

## Slide 15: Finding beam position with a lookup table
We compute the simulated V with a look-up table and 1st order approximation:
Va
Vb

## Slide 16: Numerical simulations
Va
Vb
Or with a second order approximation, which give much better results:

## Slide 17: Numerical simulations
Va
Vb
Experimental issue with the look-up table method: it doesn’t always converge!
Xn
Xn+1
Here (Xn) doesn’t converge and oscillate between
2 points

## Slide 18: Numerical simulations
Experimental issue with the look-up table method: it doesn’t always converge!
My solution to this: if (Xn) didn’t converge after some iteration (usually 5 iterations are enough), then the next Xn is the mean value of 2 Newton’s iterations.

## Slide 19: Numerical simulations
But it is not enough! In some rare cases we have this:
Xn
Xn+1
Here (Xn) oscillate between 3 points!
Xn+2
Once more a solution is: if (Xn) didn’t converge after some iterations, then the next Xn is the mean value of 3 Newton’s iterations.
 With this solution I couldn’t find any cases where (Xn) didn’t converge.

## Slide 20: Results
A few words about performance of the look-up table method:
Using a 601 x 601 table (generated in 5 minutes by G. Rehm’s script),
We can recover the position of a measurement with nanometer accuracy,
It take 3.5 ms to find a position with this method using Matlab (it is not yet faster than using the numerical solver).

## Slide 21
Introduction
Finding beam position with a numerical solver
  Method
  Numerical solvers available
Finding beam position with a lookup table
  Method
  Non-converging situations and solution
  Results
Future developments

## Slide 22: Conclusion
Using numerical method we can (in principle) get rid of non-linear effect of BPM blocks
The look-up table method gives also good results but may have some issues finding a solution
Next steps:
Use CST to generate the lookup table (is it possible to generate the table in a reasonable amount of time?)
Try to implement the lookup table method in C++ to improve execution time