## Slide 1: The AGIPD Project December 2016 – XDAC

## Slide 2: The AGIPD Read Out-Principle
Sensor
Electronics per pixel
ASIC periphery
Chip output driver
Mux
HV
+
-
THR
DAC
SW
CTRL
Analog Mem
Analog Mem
CDS
RO Amp
Calibration circuitry
Adaptive gain amplifier
352 analog memory cells
…
…
Read Out bus
Pixel matrix
2
Notes: Mention random vetoing possibility.

## Slide 3: The AGIPD Parameters
200 x 200 micron2 pixels
352 storage cells + veto possibilities.
Minumum signal ~ 300 e- =
  0.1 photon of 12.4 keV
Maximum signal ~ 33x106 e- =
  104 photons of 12.4 keV
4.5 MHz frame rate
64 x 64 pixels per ASIC
2 x 8 ASICs per module (128x512 pixels, no dead area)
4 modules per quadrant
Total of 1 Million pixels
3

## Slide 4
The AGIPD system
4

## Slide 5
Fully assembled quadrant
5

## Slide 6
The 1-Million pixel system

## Slide 7
Prepare the detector for removing a quad.  1.5 days with 2 persons
Exchange a quad.  ½ day with 2 persons
Put everything back into place, close the detector, (mount pump) and cooling tubes.  2 days with 2 persons
Calculate time for pumping down and cooling down before restart of the system.
Time to exchange a quadrant

## Slide 8
  Firmware with complete readout chain adapted for uC Integration fully operational.
  Double Images(Analog + Gains) per memory cell (AGIPD 1.0) agreed with XFEL and Train builder Format implementation updated
  Clock & Control System hardware integration is done. Full communication protocol is implemented. Relevant XFEL bunch structure and timing on firmware level are implemented and can be tuned in real time via user GUI
Status Firm- and Software

## Slide 9
  HERA South Integration tests
    As the Train Build Format test with XFEL PC layer was successful, XFEL would like to wait until the full 1M detector ready in HERA South for test with real TB
    Clock & Control System integration and tuning with HERA South environment
  Remote upgrade firmware via Ethernet
  Firmware optimization for AGIPD 1.1
Outlook Firm- and Software

## Slide 10
Power-up procedure
Partially coded to µC-firmware
Alarms are available
Temporary command
reception in µC is
available
Analogue cards:
Per default on
Most cards have own hardware protection

## Slide 11
Calibration activities
AGIPD calibration on delivery | 
pedestal maps noise maps Gain maps for 3 gain stages | Raw data, calibration routines and processed data delivered to XFEL
Raw data
Sum of 50k frames
Dead
ASIC
Pedestal corrected
Sum of 50k frames
4-5Photons
Ongoing activities
Analysis of „Grid-data“

## Slide 12
Calibration activities
Ongoing activities
Analysis of „Grid-data“ Threshold table for on the fly gain bit correction. Bad pixel maps
AGIPD calibration on delivery | 
pedestal maps noise maps Gain maps for 3 gain stages | Raw data, calibration routines and processed data delivered to XFEL

## Slide 13
Calibration activities
Ongoing activities
Analysis of „Grid-data“ Threshold table for on the fly gain bit correction. Bad pixel maps Test of calibration procedure in the course of analysis of protein crystallography data from beam line P11
Experiment, calibration and data correction and analysis were performed in close cooperation with EMBL and FS-CXI

## Slide 14
Calibration activities of the next 6 months
Calibration activities of the next 6 months
Optimize calibration and DAQ patterns for AGIPD 1.1 Modification of calibration routines for AGIPD 1.1 Systematic droop investigation for AGIPD 1.1 Calibration of AGIPD 1.1 for MID delivery and SPB replacement1

## Slide 15: Wafer-Scale Testing of AGIPD 1.1
A testing procedure for AGIPD 1.1. ASICS has been developed
It uses:
Cascade Mirotec (SUSS) PA200 probe station
Custom probe card
PSI “Chiptesterbox” (based on Mythen readout)
LABview and custom software
What it tests:
Baseline & Gain in Hi-Gain mode for
  all 352 storage cells with
  Pulsed Capacitor stimulus
The 3 gain bit levels for
  All 352 storage cells
Medium and low gain
  for one storage cell with
  Current Source stimulus
Power consumption

## Slide 16: Wafer-Scale Testing of AGIPD 1.1
Raw data: Test signal in even and odd pixels
Calculated gains and pedestals
Gain- and pedestal matrices written to disc
Cuts
Evaluation Results (from A10_cteval.vi)
Cuts
Matrices read from disk
Statistics (Pedestal and separation of gain bits)
Detailed results, e.g. dead storage cell map
The LABview software consists of four parts:
A10_wprobe.vi - A routine to step the wafer, to accquire the data and to format it into matrices written to disc. It calls the
A10_cteval.vi - Evaluation Routine for every Chip and
A10_wmap.vi at the end, which produces a graphical wafer map.
A10_eval.vi – same as A10_wprobe.vi, but reading from disk

## Slide 17
Delivery of current system to XFEL HERA-South Dec. 2016
  Support of XFEL during integration and operation
Production of new Front-end Modules (AGIPD 1.1)
  At PSI: first modules bump-bonded mid January 2017; followed by assembly and testing at DESY, end January. Followed by rest of modules in February.
  At IZM: first modules bump-bonded by mid January; followed by testing at DESY in January; followed by more modules in February.
  New mounting procedure (better thermal contact) by mid January
  Update of Firm and software in January
  Calibration in February and March (at DESY), using MID chamber
  Delivery AGIPD 1.1 Modules to Eu.XFEL April 2017. Modules for testing will be delivered much earlier (January)
Timeline until acceptance 1st 1M
17

## Slide 18
All parts are in hand (except for front-end modules)
Finish assembly of system in January.
  Note: we can start as soon as the SPB system is out of the lab
  Note: vacuum chamber plus motion stages area already commissioned
  Assembly of quadrants
  Mounting quadrants into vacuum chamber
  Mounting of all flanges onto chamber
Start commissioning system in February, using the AGIPD 1.1 modules for SPB.
Continue commissioning system until April 2017.
Project until delivery 2nd 1M
18