## Slide 1: Photon beamlines
Harald Sinn
Commissioning Workshop
Jork, 24-25.11.2016

## Slide 2: Beamline design
So far only two branches are being built

## Slide 3: Overview
TR
Imager
Filter
SRA
K-Mono
SR
Imager
XGM
Solid Attenuator
CRL1
2D Imager
M1
M2
Beam loss monitor
MCP
Pop-in monitor
Shutter XS3
Hirex
Beam loss monitor
M3
Beam loss monitor
Mono 1
Beam loss monitor
XGM
CRL2
Pop-in monitor
Shutter FXE
Pop-in monitor
Shutter SPB
Vacuum
Equipment Protection
Motion
Tunnel SASE 1
Motion
indicators of problems
switch to other screen
grey tabs switch to
sub-screens
Pop-in monitor
Pulse picker
Attenuator
Mono 2
Spectrum Analyzer

## Slide 4: Imagers
Can accept only one pulse per train (exception: attentuator in)
Moves YAG in with step motor
In EPS, MPS systems
Calibration of screen positions (through software)
LED light to see some signal (and calibration marks?)
Different types differ in size of YAG and resolution
SR-Imager has photodiode
screen
SASE1 Pop-in Monitor XTD2
0.03
0.00
LED
beam
CAM
out
in
on off
on off
(Proposed software panels are shown, real ones will be a bit different)
Responsible: Andreaas Koch/ Photon Diagnostics

## Slide 5: Imagers (Transmissive, 2D, Pop-in Monitors IMPG 1+2)
Responsible: Andreaas Koch/ Photon Diagnostics

## Slide 6: Simulation of Imager Views
from data base XPD:
Liubov Samoylova (X-ray Optics)

## Slide 7: Get FEL simulation data for commissioning: web access to FAST code (M.Yurkov, E. Schneidmiller) results
commissioning parameters
in.xfel.eu
Liubov Samoylova (X-ray Optics)

## Slide 8: Pick up the selected pulse data in hdf5 format
Structured HDF5 files contain all information about FEL simulation including input parameters and post-processing data
data available for download for 72 h
Liubov Samoylova (X-ray Optics)

## Slide 9: Propagate wavefront hdf5 data using WPG wave optics software
WPG software package on /afs/desy
Special commissioning folder, containing
  metrology data for the mirrors
  python scripts with layouts including all optical beamline components
  scripts can be run as Jupyter notebooks (interactive interface running in web browser)
  Can be used ‘real time’ by commissioning team
Liubov Samoylova (X-ray Optics)

## Slide 10: Filter

## Slide 11: Filter
out position
empty
Al
Ni
Cu
Mo
Al
beam
Responsible: Wolfgang Freund/ Photon Diagnostics
only 1 pulse/train or spontaneous radiation!
requires detector (diode in SR imager, XGM)
ideally also intensity monitor before
can calibrate photon enery to 1 eV precision

## Slide 12: Cu- Filter around 9 keV can be used for wavelength calibration
8979 eV +- 0.5 eV
source: CRXO website http://henke.lbl.gov
Nickel foil:
8333 eV
scan of undulator gap or electron energy

## Slide 13: SRA

## Slide 14: Spontaneous Radiation Aperture (SRA): The only slit in SASE1 tunnels
beam
responsible: Xiaohao Dong, X-ray Optics

## Slide 15: K-mono (‘K’ from undulator-k)

## Slide 16: K-mono
Responsible: Wolfgang Freund/ Photon Diagnostics
View inside the SASE3 unit

## Slide 17: K-mono (2011 design)

## Slide 18: K-mono in action
two undulator segments
data acquisition faster, but
requires bump in electron trajectory
scan of Kmono closing one or the other undulator segement (spont. rad)

## Slide 19: XGMs
Intensity + beam position
Only two for cost reasons
No motions
DOOCS controled
Jan Grünert (WP74) + Kai Tiedtke group (DESY)

## Slide 20: Solid Attenuators
Different design, similar functionality for instrument alignment
protects downstream components
Responsible: Harald Sinn, WP73

## Slide 21: Solid Attenuator

## Slide 22: Attenuator plates mounted
(Feb 2015 )
In beam direction (downstream) order (SASE1)

## Slide 24: Solid Attenuator panel
10-1
10-2
10-3
S5
6.3 keV
A1
A2
A3
A4
A5
A6
A7
A8
A9
3.1 10-1
Full beam
SASE1 Solid Attenuator
EPS

## Slide 25: Compound Refractive Lenses (CRLs)
Collimation + weak focusing (similar M2 bender)
tight focusing, backup for SPB KB mirrors

## Slide 26: Compound Refractive Lenses (CRLs)
n=1
n=0.999999
about 100 m !
Equivalent: Optical diverging lens
Beryllium

## Slide 27: CRLs in XTD2+9
Entire vessel 5-dim adjustable
Each arm x-y adjustabe
B4C apertures on each arm
Thermometers
Can be damaged > 60 pulses
Can damage beam pipes + shutters
Not needed in phase 1
Lewis Batchelor, CIE

## Slide 28: Mirrors

## Slide 29: Mirror M1 + M2 (offset mirrors)
Tungsten beam stop
M1+M2 are photon beam chicane for safety reasons
Chicane is defined by collimators (ColA1: 12.5x20 mm before M1)
Variable incidence angle used to adjust to varying beam size
Incidence angle 1.1 mrad .. 3.6 mrad
Offset 25 mm .. 85 mm
Beam is steered back over tunnel length

## Slide 30: Mirror M3 (distribution mirror)
Incidence angle 1.35 mrad defined by distance to tunnel end
‘No-clipping’ condition at M3 achieved by intermediate focus (in particular needed for low energies)
M2 has therefore adjustabe radius (∞ - 50 km)
Only 1D intermediate focus (much less damaging than CRLs)
With constant intermediate focus position same beamsize as natural beam in experiment hall
430 m

## Slide 31: Mirrors: Technical realisation
Bending mechanism prototype

## Slide 32: Temperature drift of bender
Capacitive Sensors
Fizeau Interferometer
Temperature
capacitive sensors
Stability on 10 nm level during
0.3 °C drift in mechanical bender !

## Slide 33: Mirror chambers
No translation slides, probably close to 10 nrad reproducibility, but:
Parasitic motion: ≈10 µrad /mm
However: 1 mm translation of M2 changes 40 µrad reflection angle anyway
Responsible: Antje Trapp, X-ray Optics

## Slide 34: Mirror (‘CHOMs’) panels (= Chamber for horizontal offset mirrors)
Responsible: Antje Trapp, X-ray Optics

## Slide 35: Caution: Liquid Metal Cooling!

## Slide 36: Photon beam loss monitor
responsible: Natalia Gerasimova, X-ray optics

## Slide 37: Fast shutdown of beam: The Photon Beam Loss Monitor concept
Connects directly to E-machine MPS, shuts down beam after 30-50 pulses
MPS
responsible: Natalia Gerasimova, X-ray optics

## Slide 38: PBLM after every beam steering element
after horizontal mirrors
at the end of beamlines
responsible: Natalia Gerasimova, X-ray optics

## Slide 39: PBLM panel
mechanical motions + cameras in Karabo
PMT (photomultiplier) readout and seetings through DOOCS
direct connections to MPS
responsible: Natalia Gerasimova, X-ray optics

## Slide 40: MCP
MCPs + Diodes for SASE search and optimization
Can work in direct beam or in beam through M1+M2
Single shot sensitive!
responsible: Jan Grünert, X-ray diagnostics

## Slide 41: Hirex
responsible: Naresh Kujala, X-ray diagnostics

## Slide 42: Hirex spectrometer
Can provide spectrum from single pulse measurement
Semi-transparent
Can stand 1-50 pulses/train
Useful for laser tuning (shows bandwidth of X-rays)
responsible: Naresh Kujala, X-ray diagnostics

## Slide 43: Hirex spectrometer
responsible: Naresh Kujala, X-ray diagnostics

## Slide 44: 500 m empty pipe

## Slide 45: 500 m empty pipe

## Slide 46
upstream collimator (Shutter in XS3)
SASE beam
beampipe (100 mm diameter)
end of pipe

## Slide 47: Alignment of beam transport
13 m
113 m
5 m
530 m
1.4 m
screens
mirrors
Alignment accuracy philosophy:
1 mm / 5 m = 100 mm / 500 m
SASE1

## Slide 48: Can we damage anything during first photon beam commissioning ?
B4C shutters etc
melting of beam pipes
Yes, we can!

## Slide 49: Further help: Ray tracing plots (Alfresco WP73)

## Slide 50: Mono 1+2
responsible: Xiaohao Dong, X-ray Optics

## Slide 51: Monochromator

## Slide 52: Mono 1
Ty
SASE1 Mono 1
Piezo Pitch
Pico Pitch
ACCM_pitch
Pico Roll
Tx
T2: 107.3 K
T1: 105.4 K
T0: 95.4 K
Heater 0
Heater 1
Heater 2
Energy
Scan
8.998 keV
responsible: Xiaohao Dong + Liuba Samoylova X-ray Optics

## Slide 53: Mono 2
Ty
SASE1 Mono 2
Piezo Pitch
Pico Pitch
ACCM_pitch
Pico Roll
Tx
T2: 107.3 K
T1: 105.4 K
T0: 95.4 K
Heater 0
Heater 1
Heater 2
Energy
Scan
8.998 keV

## Slide 54: SPB devices (CRL, Pulse picker, Attenuator)

## Slide 55: Pulse picker
10 Hz operation
B4C disk + Tantalum
Should be able to survive 60 pulses, but rotor gets warm over time.
Will be tied to MPS to reduce beam mode
Responsible: Lewis Batchelor, CIE

## Slide 56: Attenuator (SPB)
Materials: Diamond, B4C, Silicon
Double thickness when going to next arm
Attention: Cu+Fe+H2O crosses beam! (need interlock based on encoder at some point)

## Slide 57: Shutters
Responsible: Fan Yang , X-ray Optics

## Slide 58: Shutter
Fan Yang/X-ray Optics
Will be controlled primarily through DESY MPS group software
Can stand up to about 300 pulses/train (Mode 3)
Power absorbers always closes before tungsten shutter
Can be damaged by CRL focus

## Slide 59
Finish!