## Slide 1: Clarification of the term ‘Operational’
‘Operational’ means that a device can be operated from a control room with/for x-ray beam or other final functionality. Basic functionality allowing the main purpose of the device shall be enabled. For this, at a minimum, the following steps have to be completed.
  Mechanical installation
    installation of all hardware,
    pre-alignment (without beam),
    connection to adjacent devices,
    establish vacuum conditions (where applicable)
  Cabling, control electronics hardware
    All cables installed and wired both sides,
    Control (& read-out) electronics hardware installed and documented (Eplan, WS-CAD)
  Interface electronics
    Electronic hardware locally tested,
    Beckhoff fw, micro-TCA fw (or other low level servers) installed and tested
    Inclusion in Beckhoff loop (or network)
  Controls software
    High-level controls and read-out software installed
    Basic functionality installed and tested
    Handover to ‘device-owner’
  Technical commissioning
    Test of (basic) device functions
    Remote operation tested over several days