## Slide 1: WP74 status w3 (20.1.2017) – past week progress
XGM @ XTD2 connected to vacuum beamline on both sides
XGM @ XTD9 re-aligned (survey&alignment)
First time achieved (karabo-)control of an MPOD-crate IN TUNNEL (at MCP in XTD2)
Local test of Beckhoff/Motors at MCP@XTD2 done, but issues
Manual alignment of camera box at MCP@XTD2 done
HIREX@XTD9 connected to vacuum beamline (ALL interfaces)
Cabling in progress for XGM@XTD2: vacuum  Beckhoff
Cabling in progress for K-mono + SR-imager: device  crate/rack cables installation
Started implementation of trigger supply from MTCA  cameras etc. ( AE)
New group member: introduction of Theophilos Maltezopoulos for gas-based physics (contact for XGM)

## Slide 2: Issues
During local test of motors at MCP, two main issues were noted:
  motor move too slow
  Port 1 manipulator blocks when driven into end switch
Trigger hardware (cables+adapter boxes) required from AE, must be installed
DAQ software was not ready to acquire data from Struck ADC boards (e.g. at MCP)
Our CAS USP (W.Ehsan) cannot enter the tunnel (neither physically nor computerally)
XGM@XTD2 is mechanically-vacuum-connected, but not the control system !!!
“bring x-ray beam to end of XTD9” implies “First Lasing” in XTD2 !!

## Slide 3: Clarification of the term ‘Operational’
‘Operational’ means that a device can be operated from a control room with/for x-ray beam or other final functionality. Basic functionality allowing the main purpose of the device shall be enabled. For this, at a minimum, the following steps have to be completed.
  Mechanical installation includes
    installation of all hardware,
    pre-alignment (without beam),
    connection to adjacent devices,
    establish vacuum conditions (where applicable)
  Cabling, control electronics includes
    All cables installed and wired both sides,
    Control (& read-out) electronics hardware installed,
    Electronic hardware locally tested,
    Beckhoff fw, micro-TCA fw (or other low level servers) installed and tested
  Controls software
    High-level controls and read-out software installed
    Basic functionality installed and tested
    Handover to ‘device-owner’
  Technical commissioning
    Test of (basic) device functions
    Proven retractability
    Remote operation tested over several days