## Slide 1: XGM – progress week 8 / 2017
Dr. Jan Grünert
X-ray Photon Diagnostics
Group Leader
Hamburg, February 24th, 2017
Contains information from WP74 team
and from K.Tiedtke / F.Jastrow from DESY-FS-FL-D

## Slide 2: XGM - Progress (w8)
XGM@XTD2
  local Beckhoff test, including Gas-Supply: DONE (some minor rework)
  Pressure (22.2.): 1.24E-7 mbar (differential pumping is still off)
  HV test on XGM@XTD2: delayed to next Monday
  Survey&Alignment: DONE (acceptable change of 0.05 mm  < 0.15mrad change about y-axis)
XGM@XTD9
  Repair (exchange by DESY of resistors in both XGMD-chambers): DONE
  Pump-down after repair  9E-6mbar (Thursday 16h30h)
  Leak check: today (Friday 24.2. / now)
Trigger cabling (rack-to-devices) in XTD2+XTD9 (~2-3 days) COMPLETED
Grounding : check actual ground connections XTD2+9+10 with experts (today Friday 24.2. / now)

## Slide 3: XGM – Repair Progress Summary (w8)
XGM | Preparation | Resistor exchange | Leak test | HV test | Survey
HERA XGM01 | done | done | done |  | 
XTD2 XGM02 | done | done | done |  | done
XTD9 XGM04 | done | done |  |  | 
XTD10 XGM07 |  |  |  |  | 
XHE3 XGM05 |  |  |  |  | 
XHE3 XGM06 |  |  |  |  | 

## Slide 4: XGM - Plan week 9
Planning of Repair XGM@XTD10
XGM@XTD9: wiring of vacuum-to-Beckhoff
Continue tests of serial communication from DOOCS to Keithleys/SRG/RVC300/HV-crate

## Slide 5: WP74 – other devices SASE1– w8
HIREX: Beckhoff crates + cables installed  ready for local Beckhoff test
SR imager: installed full spring compensation