## Slide 1: WG1 Cavity Performance from FRIB
K. Saito MSU/FRIB SRF Development Manager

## Slide 2: FRIB Production Cavities
FRIB certified cavities so far:
16 b=0.041 QWRs, 96 b=0.085 QWRs, 72 b=0.29 HWRs,
and 81 b=0.53 HWRs, total 268/322 (81%).
These statistics can give a valuable information about cavity performance limitation, which suggests needed R&Ds for future high gradient (> 10MV/m) low/medium beta cavities.
FRIB cavity basic information:
  Fabricated high purity Nb material with RRR > 250
  Bulk etching by BCP (> 120mm),
  600OC - 10 hrs hydrogen degassing,
  Light etching by BCP( ~10mm),
  HPR with ultrapure water,
  No low temperature bake, no impact on 2K cavity performance
  Certification test at 2K.

## Slide 3: FRIB Cavity Performance (VTA)
0.041 QWR (16)
0.085 QWR (86)
0.29 HWR (73)
0.53 HWR (61)

## Slide 4: Statistics FRIB Cavity Performance Limit
 | X-ay | Thermal Breakdown | Low LHe level | RF Power | Administration limit | Eacc, max [MV/m] | QO at operation gradient
0.041QWR (16) | 0% (0) | 19% (3) | 25% (4) | 0%(0) | 6% (1) | 10.5 ± 1.3 (Spec 5.1) | (5.7 ± 1.2) x 109 (Spec 1.4 x 109)
0.085QWR (86) | 3% (3) | 8% (8) | 13% (13) | 4% (4) | 71% (69) | 9.1 ± 0.9 (Spec 6.7) | ( 3.9 ± 1.4) x 109 (Spec 2.0 x 109)
0.29 HWRs (72) | 1% (1) | 35% (25) | 11% (8) | 0% (0) | 53% (38) | 12.6 ± 1.4 (Spec 9.2) | (1.4 ± 0.4) x 1010 (Spec 6.7 x109)
0.53 HWRs (84) | 0% (0) | 18% (15) | 20% (17) | 13% (11) | 49 % (41) | 11.6 ± 0.9 (Spec 8.9) | (1.9 ± 0.7) x 1010 (Spec 9.2 x 109)
More than half of FRIB cavities has excellent performance (Administration limit). 0.041QWR data base is not yet detail analyzed but all cavities are good.
X-ray limit (500mR/hr) is very less thanks to HPR.
Thermal breakdown occurs with a rather high probability ~ 20%. We suspect EBW defects rather than material due to the complicated welding configuration of low/medium beta cavities. We observe “Cats eye” defect on EBW seam in some cases
Excellent performance

## Slide 5: Cavity Performance Dependence on Ep and Bp
Cavity performance dependence on Ep and Bp, for examples FRIB b=0.041 QWRs (16) and b=0.29 HWRs (72).
Several FE results are mixed.
Observes Q-drop in both families, and 0.085QRWs and 0.53HWRs also same
Q-drop on-set
Q-drop on-set
0.041QWR
0.29HWR
K. Saito , June 26 2018, TTC meeting RIKEN WG1
, Slide 5

## Slide 6: High Field Q-Slope
Pure high field Q-slop (HFQS) without X-ray is often observed in every FRIB cavity family.
0.041QWR
0.29HWR
0.085QWR
0.53HWR
0.041QWR-Xray
0.085QWR-Xray
0.29HWR-Xray
0.53HWR-Xray
Eacc ~ 8 MV/m
Eacc ~ 7 MV/m
Eacc ~ 9.5 MV/m
Eacc ~ 11 MV/m
No X-ray
No X-ray
No X-ray
No X-ray
HFQS
HFQS
HFQS
HFQS

## Slide 7: Q-Drop and The Onset Field
Observed Q-drop in every cavity family, of which potential cause is field emission (FE) or high field Q-slope (HFQS).
FE is often observed but HFQS is also seen with a high probability in FRIB cavities.
The onset field is coincident at a Bp ~ 84 mT in every cavity family, which corresponds to Eacc=20MV/m for ILC cavity, which is known the onset filed of HFQS with BCP’ed cavities.
FE and HFQS must be improved for high gradient operation > 10 MV/m future low/medium beta cavities.
 | Field Emission (X-ray) | High Field Q-Solpe (No X-ray)
0.041QWR | 47% (7) | 53% (8)
0.085QWR | 77% (74) | 23% (22)
0.29HWR | 43% (30) | 57% (40)
0.53HWR | 80% (60) | 20% (15)
 | Q-drop onset Eacc [MV/] | Q-drop onset Ep [MV/m] | Q-drop onset Bp [mT]
0.041QWR | 6 | 48 | 85
0.085QWR | 7 | 41 | 84
0.29HWR | 11 | 47 | 85
0.53HWR | 9.5 | 34 | 83

## Slide 8: X-ray Events during Cavity Testing at 4K and 2K
X-ray during MP conditioning and some discreet X-ray event level are observed.
RF processing of the X-ray event levels sometimes produce FE.
X-ray by the gap MP (very low field) is not observed in the graph due to the sensitivity of the detector but which is almost observed in the Qo vs Eacc curve for every cavity families.
Example
0.53HWR

## Slide 9: X-ray Event and FE Onset Field with HWRs
Chart: 0.29HWR X-ray events and FE onset
| 0 | 0.2 | 0.4 | 0.6 | 0.8 | 1 | 1.2 | 1.4 | 1.6 | 1.8 | 2 | 2.2000000000000002 | 2.4 | 2.6 | 2.8 | 3 | 3.2 | 3.4 | 3.6 | 3.8 | 4 | 4.2 | 4.4000000000000004 | 4.5999999999999996 | 4.8 | 5 | 5.2 | 5.4 | 5.6 | 5.8 | 6 | 6.2 | 6.4 | 6.6 | 6.8 | 7 | 7.2 | 7.4 | 7.6 | 7.8 | 8 | 8.1999999999999993 | 8.4 | 8.6 | 8.8000000000000007 | 9 | 9.1999999999999993 | 9.4 | 9.6 | 9.8000000000000007 | 10 | 10.5 | 11 | 11.5 | 12 | 12.2 | 12.4 | 12.6 | 12.8 | 13
 | 0.9 | 0.3 | 0 | 0.042857 | 0.014286 | 0.042857 | 0.057143 | 0.042857 | 0.028571 | 0.028571 | 0.071429 | 0.028571 | 0.085714 | 0.18571 | 0.47143 | 0.77143 | 0.84286 | 0.94286 | 0.95714 | 0.94286 | 0.94286 | 0.87143 | 0.7 | 0.41429 | 0.11429 | 0.11429 | 0.11429 | 0.11429 | 0.1 | 0.18571 | 0.25714 | 0.21429 | 0.21429 | 0.12857 | 0.15714 | 0.11429 | 0.028571 | 0.014286 | 0.028571 | 0.071429 | 0.1 | 0.057143 | 0.042857 | 0.014286 | 0.014286 | 0.042857 | 0.028571 | 0.028571 | 0.028571 | 0.028571 | 0.028571 | 0.014286 | 0.028571 | 0.014286 | 0.014286 | 0.014286 | 0.014286 | 0.014286 | 0.014286 | 0.014286
 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.042857 | 0.014286 | 0 | 0.085714 | 0.028571 | 0.014286 | 0.014286 | 0 | 0.042857 | 0.028571 | 0.014286 | 0.014286 | 0 | 0.014286 | 0.028571 | 0.014286 | 0 | 0 | 0.014286 | 0 | 0 | 0 | 0 | 0.014286 | 0 | 0.014286 | 0 | 0 | 0 | 0 | 0 | 0 | 0.014286
Chart: 0.53HWR X-ray events and FE onset
| 0 | 0.2 | 0.4 | 0.6 | 0.8 | 1 | 1.2 | 1.4 | 1.6 | 1.8 | 2 | 2.2000000000000002 | 2.4 | 2.6 | 2.8 | 3 | 3.2 | 3.4 | 3.6 | 3.8 | 4 | 4.2 | 4.4000000000000004 | 4.5999999999999996 | 4.8 | 5 | 5.2 | 5.4 | 5.6 | 5.8 | 6 | 6.2 | 6.4 | 6.6 | 6.8 | 7 | 7.2 | 7.4 | 7.6 | 7.8 | 8 | 8.1999999999999993 | 8.4 | 8.6 | 8.8000000000000007 | 9 | 9.1999999999999993 | 9.4 | 9.6 | 9.8000000000000007 | 10 | 10.5 | 11 | 11.5 | 12
 | 0 | 0.81579 | 0.013158 | 0 | 0 | 0 | 0.013158 | 0.013158 | 0 | 0.013158 | 0.052632 | 0.59211 | 0.80263 | 0.84211 | 0.88158 | 0.90789 | 0.94737 | 0.94737 | 0.84211 | 0.34211 | 0.19737 | 0.68421 | 0.90789 | 0.77632 | 0.69737 | 0.51316 | 0.39474 | 0.32895 | 0.19737 | 0.065789 | 0.013158 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.013158 | 0.026316
 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.013158 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.026316 | 0 | 0 | 0.026316 | 0.052632 | 0.092105 | 0 | 0.065789 | 0.052632 | 0.039474 | 0.013158 | 0.013158 | 0.10526 | 0.039474 | 0.039474 | 0.065789 | 0.026316 | 0.013158 | 0 | 0.026316 | 0 | 0 | 0 | 0 | 0.013158 | 0.026316 | 0 | 0 | 0 | 0.013158
Gap MP
2P-1st MP at short
MP at FPC port
MP + ?
FE onset field
Gap MP, 2-point 1st order MP at short area, and MP at FPC port area are observed.
FE onset is correlated to the discreet X-ray event levels on 0.29HWRs but not on 0.53HWRs
Need to understand the mechanism of the discreet X-ray event levels to eliminate FE.
2P-1st MP at short
Gap MP
FE onset field

## Slide 10: X-ray Event and FE Onset Field with QWRs
Chart: 0.041QWR X-ray events and FE onset
| 0 | 0.2 | 0.4 | 0.6 | 0.8 | 1 | 1.2 | 1.4 | 1.6 | 1.8 | 2 | 2.2000000000000002 | 2.4 | 2.6 | 2.8 | 3 | 3.2 | 3.4 | 3.6 | 3.8 | 4 | 4.2 | 4.4000000000000004 | 4.5999999999999996 | 4.8 | 5 | 5.2 | 5.4 | 5.6 | 5.8 | 6 | 6.2 | 6.4 | 6.6 | 6.8 | 7 | 7.2 | 7.4 | 7.6 | 7.8 | 8 | 8.1999999999999993 | 8.4 | 8.6 | 8.8000000000000007 | 9 | 9.1999999999999993 | 9.4 | 9.6 | 9.8000000000000007 | 10 | 10.5 | 11 | 11.5 | 12
 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.066667 | 0 | 0 | 0 | 0 | 0 | 0.066667 | 0 | 0 | 0.13333 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.066667 | 0 | 0 | 0 | 0.066667 | 0 | 0 | 0 | 0 | 0 | 0
 | 0.13333 | 0 | 0 | 0.26667 | 0.26667 | 0.13333 | 0.066667 | 0.066667 | 0.2 | 0.2 | 0.13333 | 0.13333 | 0.13333 | 0.066667 | 0.066667 | 0.13333 | 0.2 | 0.2 | 0.2 | 0.13333 | 0.2 | 0.26667 | 0.2 | 0.13333 | 0.13333 | 0.13333 | 0.13333 | 0.066667 | 0.13333 | 0.13333 | 0.13333 | 0 | 0.066667 | 0 | 0 | 0.066667 | 0 | 0 | 0 | 0.066667 | 0 | 0 | 0.13333 | 0 | 0.13333 | 0 | 0.13333 | 0.066667 | 0.066667 | 0.13333 | 0.066667 | 0.066667 | 0.066667 | 0 | 0
Chart: 0.085QWR X-ray events and FE onset
| 0 | 0.2 | 0.4 | 0.6 | 0.8 | 1 | 1.2 | 1.4 | 1.6 | 1.8 | 2 | 2.2000000000000002 | 2.4 | 2.6 | 2.8 | 3 | 3.2 | 3.4 | 3.6 | 3.8 | 4 | 4.2 | 4.4000000000000004 | 4.5999999999999996 | 4.8 | 5 | 5.2 | 5.4 | 5.6 | 5.8 | 6 | 6.2 | 6.4 | 6.6 | 6.8 | 7 | 7.2 | 7.4 | 7.6 | 7.8 | 8 | 8.1999999999999993 | 8.4 | 8.6 | 8.8000000000000007 | 9 | 9.1999999999999993 | 9.4 | 9.6 | 9.8000000000000007 | 10 | 10.5 | 11 | 11.5 | 12
 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.010417 | 0.020833 | 0 | 0.020833 | 0.041667 | 0.020833 | 0.052083 | 0.03125 | 0.010417 | 0.010417 | 0.010417 | 0.03125 | 0.052083 | 0.072917 | 0.03125 | 0.020833 | 0.041667 | 0.052083 | 0.041667 | 0.03125 | 0 | 0.052083 | 0.020833 | 0.03125 | 0.010417 | 0.010417 | 0.03125 | 0 | 0 | 0 | 0 | 0 | 0 | 0.020833 | 0 | 0 | 0 | 0 | 0
 | 0.76042 | 0.020833 | 0.010417 | 0.80208 | 0.083333 | 0.010417 | 0 | 0 | 0.010417 | 0.020833 | 0.010417 | 0 | 0.020833 | 0.03125 | 0.010417 | 0.020833 | 0.020833 | 0 | 0.010417 | 0 | 0 | 0 | 0.010417 | 0.010417 | 0.010417 | 0.010417 | 0 | 0 | 0 | 0.020833 | 0.010417 | 0.010417 | 0.03125 | 0.041667 | 0.03125 | 0.0625 | 0.072917 | 0.083333 | 0.0625 | 0.072917 | 0.052083 | 0.041667 | 0.052083 | 0.03125 | 0 | 0 | 0 | 0 | 0 | 0 | 0.010417 | 0 | 0 | 0 | 0
Gap MP, 2-point 1st order MP at short area are observed.
FE onset is correlated to the discreet X-ray event levels on both 0.041 and 0.085QWRs.

## Slide 11: Summary
FRIB production cavity statics shows that still field emission and thermal breakdown are concerned as the major performance limitation for Eacc > 10MV/m.
EBW defects are suspected for the thermal breakdown because of the complicated welding scheme of low/medium beta cavities.
In addition, HFQS is another concern with BCP’ed cavities.
The multipacting does not give the performance limitation but requires time consuming RF conditioning.
Rinse ports like FPC one needs to more careful rinsing to reduce MP.
The discrete X-ray event levels some times produce FE. This mechanism must be understood for future high gradient operation low/medium beta cavities.

## Slide 12: Low Temperature Baking Benefit on 4K/2K Performance on ReA3 0.085QWRs
ReA3 SC252
Low temperature (LT) bake was investigated on ReA3 0.085QWRs prior to the FRIB cavity production.
  Observed a benefit at 4K but no benefit at 2K
LT Baking impacts Qo at 4K which BCS Surface resistance is dominant but less benefit at 2K due to very small RBCS (80.5MHz).
ReA3 is operated at 4K and employed LT bake.
FRIB is operated at 2K, LT bake was not employed to save time.
4K
2K
X-ray at 2K
ReA3 cavities 4K

## Slide 13: Multipacting (MP)
MP dese not limit the performance but is painful in the VTA test process.
Rule of thumb or simulation.
Dedicated MP simulation at SLAC with ACE3P (LINAC2010,Tsukuba)
~
One-point n-th order (short area)
Two-point n-th order (short area)
~
Rule of thumb
MP in short gap (beam pipe)
= 2
2-poit first order MP at short area
MP at FPC port

## Slide 14: Expected MP Barrier
MP level in Eacc | 0.041QWR | 0.085QWR | 0.29HWR | 0.53HWR
Gap 1st [MV/m] | 0.035 | 0.072 | 0.98 | 1.8
One-point 1st [MV/m], Short area | 0.23 | 0.2 | 1.2 | 1.1
Two-point 1st [MV/m], Short area | 0.45 | 0.39 | 2.5 | 2.3
0.53HWR | 3D simulation
Tow point first at short area | 1.6 - 2.2 MV/m
Rinse port | 0.77 MV/m
FPC port | 3.0 - 4.1 MV/m
Rule of Thumb
3D Simulation for FRIB 0.53HWR
by ACE3P at SLAC (LINAC2010, Tsukuba), the shape is little bit different from FRIB production cavity)

## Slide 15: WG1
WG-1; What are the performance limits for low beta resonators?
There are a variety of significant low to medium beta projects being pursued using non-elliptical TEM mode cavities (QWR, HWR, single and multiple spokes). Projects include FRIB, ESS, RISP, C-ADS and RIKEN. In principal the performance limits for non-elliptical cavities should be similar to elliptical cavities but the cavities typically operate in a lower frequency range and with unique geometries. Comparing performance across the variants is not as evolved as with 1.3GHz elliptical cavities. The scope of this working group is to discuss the known limitations of TEM mode accelerating cavities. The range of discussion should include:
· Performance results Q,Rs(B) vs Ea,B from various labs
· Experience with heat treatments including doping
· Role of rf frequency
· Multipacting issues
· 4K vs 2K operation