## Slide 1
N-doping study at KEK
Temperature dependence of Nitrogen doping effects
Takafumi Okada (SOKENDAI),
Kensei Umemori, Yoichiro Hori, Eiji Kako, Taro Konomi, Hiroshi Sakai, Takeshi Dohmae (KEK),
Jyunichiro Kamiya, Shunta Kurosawa, Kenichi Takeishi (JAEA)
TTC meeting at RIKEN 2018/June/27

## Slide 2: Motivation
Nitrogen doping improve the Q value of the SRF cavity.
At KEK, Nitrogen doping for 1.3 GHz single-cell and 3-cell cavities at J-PARC furnace is success.
Deep doping for SIngle-cell cavity
1.5 times Q0 at 14 MV/m, 2K
Light doping for 3-cell cavity
1.5 times Q0 at 20 MV/m, 2K
We focus on
effects of Nitrogen doping
at other Temperature.

## Slide 3: Nitrogen doping at KEK using furnace in J-PARC
Using oil-free furnace at J-PARC (cryo-pump and three TMP. ~1×10-6 Pa)
N doping for single cell and 3-cell Nb cavities.
  Single cell (deep N-dope) :
    800℃ 20 min with Nitrogen 2.7 Pa,
    800℃ 30 min, 15 μm EP
  3-cell (light N-dope) :
    800℃ 2 min with Nitrogen 3.3 Pa,
    800℃ 6 min, 5 μm EP

## Slide 4: Test result of cavity performance
Deep doping for single-cell cavity
Cavity performance improve at under 2K.
However, unchanged Q0 value for 3-cell results at 4.2K.
2.0K
1.8K
1.6K
1.5K
BCS resistance
Light doping for 3-cell cavity
4.2K
2.0K
1.8K
1.5K
BCS resistance
VT1, VT3

## Slide 5: 3-cell cavity parameter
Tesla-like shape
Cell material is Nb from CBMM, high Ta, with 2 times melted (RRR 250-300)
Beam tube material is CBMM Nb without melting process
For 120℃ baking, EP 50 μm,
limited Eacc = 35 MV/m
VT1:Refence test limited Eacc = 35 MV/m
VT2, VT3:N-doping without magnetic field cancelling
VT4:Canceled magnetic field, with thermal gradient
We investigated the temperature dependence of the Nitrogen doping effect using 3-cell cavity

## Slide 6: Test result of cavity performance at various Temperature
We tested 3-cell cavity performance at each temperature.
  Input coupler is variable
  Cancel magnetic field by solenoid coil
  Cooling with thermal gradient by heater
The Q0 factor decrease with gradient at temperature from 4.2K to about 2.3K.
The BCS resistance obtained by Resistance-Temp fitting curve
Field emission
4.2K
3.6K
2.8K
2.3K
1.9K
2.0K
1.7K
1.6K
1.5K
1.8K
VT4

## Slide 7: BCS resistance at various Temperature
Calculate the BCS resistance from
The BCS resistance decrease with gradient at Temperature until about point.
Anti Q-slope
4.2K
3.6K
2.8K
2.3K
1.9K
2.0K
1.7K
1.6K
1.5K
1.8K
preliminary

## Slide 8: BCS resistance before Nitrogen doping
4.2K
2.0K
1.5K
1.8K
Before Nitrogen doping (Standard recipe)
We measured Q0 data at only 4.2K, 2.0K, 1.8K, 1.5K
The BCS resistance at 4.2K to 1.5K calculate by fitting function
VT1

## Slide 9: BCS resistance before Nitrogen doping
4.2K
3.6K
2.8K
2.3K
1.9K
2.0K
1.7K
1.6K
1.5K
1.8K
Before Nitrogen doping (Standard recipe)
We measured Q0 data at only 4.2K, 2.0K, 1.8K, 1.5K
The BCS resistance at 4.2K to 1.5K calculate by fitting function
The decrease of BCS resistance was not observed.
preliminary

## Slide 10: Temperature depence of BCS resistance
Normalized the BCS resistance by value at 2 MV/m
4.2K
3.6K
2.8K
2.3K
1.9K
2.0K
1.7K
1.6K
1.5K
1.8K
Light doping for 3-cell cavity
Before doping
The sign of slope is negative under 2K
The sign of slope is always positive
preliminary
preliminary

## Slide 11: Summary
KEK studied Nitrogen doping for 1.3 GHz single-cell and 1.3 GHz 3-cell cavities using J-PARC furnace.
    The single-cell for deep doping, The 3-cell for light doping
Both cavities improved Q-factor at 2K. The BCS resistance decreasing with accelerating field
The effect of Nitrogen doping was not observed at 4.2K
The effect of Nitrogen doping to BCS resistance increase with decreasing temperature
Questions
Why is the Nitrogen doping effective at temperature under 2K?
Is it correct that nitrogen doping is not effective at 4.2K?
What is the behind physics of temperature dependence of Nitrogen doping?

## Slide 12: Cavity performance tests set up
Single-cell (long N-dope)
 | EP | Cancelling magnetic fied | Cooling with thermal gradient
First Vertical test | 15 μm | ○ | ○
Second vertical test | +10 μm (total 25 μm) | ○ | ○
3-cell (short N-dope)
 | EP | Cancelling magnetic fied | Cooling with thermal gradient
First Vertical test | 5 μm | × | ○
Second vertical test | 0 μm (total 5 μm) | × | ×
Third vertical test | 0 μm (total 5 μm) | ○ | ○
Reduced magnetic field using solenoid coil
Cooling with thermal gradient using heater on upper beam pipe
Solenoid coil
Heater

## Slide 13: Temperature depence of BCS resistance
Normalized the BCS resistance by value at 2 MV/m
Ratio between Normalized the BCS resistance of N doped cavity and Normalized the BCS resistance of standard recipe cavity
4.2K
3.6K
2.8K
2.3K
1.9K
2.0K
1.7K
1.6K
1.5K
1.8K
preliminary

## Slide 14: Temperature depence of BCS resistance for Nitrogen infusion
1.9K
4.2K
3.6K
3.0K
2.2K
2.5K
2.0K
1.6K
1.8K
1.6K
1.8K
1.9K
2.0K
2.2K
2.5K
3.0K
3.6K
4.2K
preliminary

## Slide 15: Temperature depence of BCS resistance for Nitrogen infusion
1.9K
4.2K
3.6K
3.0K
2.2K
2.5K
2.0K
1.6K
1.8K
preliminary
preliminary

## Slide 16: Temperature depence of BCS resistance for Nitrogen infusion
1.9K
4.2K
3.6K
3.0K
2.2K
2.5K
2.0K
1.6K
1.8K
preliminary
preliminary

## Slide 17: BCS resistance and Residual resistance at 2K
Single-cell (deep N-dope)
The BCS resistance decreasing with gradient
Reduced the residual resistance
3-cell (short N-dope)
The BCS resistance decreasing with gradient
Reduced the residual resistance
The BCS resistance of both cavities
decreasing with gradient (Anti Q-slope) at 2K

## Slide 18: BCS resistance at 4.2K
Single-cell (deep N-dope)
3-cell (short N-dope)
Not decreasing BCS resistance with gradient (Anti Q-slope) at 4.2K