## Slide 1: Photon Run Coordinator
Magnet power supply problem: During last weekend there were many failures of the some magnet power supplies. This effectively prevented FXE/SPB to collect data. On Monday the AC was repaired, but still one failure happend on Monday noon. Tuesday/Wednesday several hardware parts and software were replaced and updated. No ‘smoking gun’ was found, but since then this failure did not happen again.
Power shutdown: A 0.6 second general power failure in Hamburg/Schenefeld area led to a complete shutdown of the Linac and cryosystems on Thursday 7am. Restart of the cryo-system happend during the day and re-tuning of the Linac+SASE during the night. After midnight SASE was established, however, the slow signal of the XGM does not give meaninful data. Theo and Fini were working on it, however they will need ZZ access this morning to XTD2 in order to fix the problem.

## Slide 2: SPB/SFX
we would like new loop projects including coordinated motion to be deployed as close as possible to the beginning of the shutdown ~18th/19th
we would like to deploy Karabo middle layers for coordinated motion on same time scale (hopefully having been *tested* with test hardware)
a number of other devices will have to be used for the first time for the mirror commissioning (these will be common to SPB/SCS/SQS systems) and not all are tested. Likely to need CAS/AE support for this
pre optic installation motion tests are scheduled for 2,3,4 July. Will contact people form other interested groups to join
it seems like the TS planned activities fit ok into our shutdown plans if dates are respected
it seems like D3 activities fit ok into our shutdown plans if dates are respected

## Slide 3: FXE
recovering from power glitch
as soon as pp laser is available again adapt laser parameters in D.05
for upcoming beamtime
finishing preparations for #002052 and starting night shift today
continuing user operation with ad-hoc procedures for laser operation
  shutter 1LSS.06 currently not available, thus an administrative procedure to safely operate a laser in D.05 had been agreed upon by the involved parties
FXE extensively used the scan tool throughout the last two weeks ->
smooth operation and very helpful, but development should not be stopped.

## Slide 4: Laser
SASE 1:
Laser shutter problems in SASE 1 are being addressed. Exact reasons for problems are still not completely known. Manufacturer is involved, we are making progress.
FXE Operation was done and will be done via "ad hoc" safety procedure until the cause of the error has been found and a solution implemented.
User operation was running according to plan (with some minor adaptations to accommodate users wishes) until 7:05 am Thursday, 07. June. Major power glitch shut down complete PP-laser and control system. At this stage (Thursday afternoon) we are still trying to re-start the laser. This is taking time due to multiple problems (network, loss of PLC settings, seeder will not start mode-locking, etc.)
IF all sub-sytems will be fine, no damage due to the hard shut-down has been incurred and the laser is up and running, there will be around one day of thermalization required.

## Slide 5: Laser
SASE 1:
Network, Karabo and PLC are back
seeder and sync back
FE started and is thermalizing overnight
high power amps will start up tomorrow
concerning laser shutter in FXE ILH, an operative shutter has already been installed yesterday (Wed), however, we are keeping up the "ad hoc" in favor of testing the shutter in this location.

## Slide 6: Vacuum
everything up and running,
no impact on the vacuum system due to the power failure on Thursday

## Slide 7: Photon Diagnostics
XPD did a group retreat 6.6.-7.6.2018 with all group members.
Power outage on 7.6. apparently didn't do any harm to diagnostics in tunnels (checked remotly)
Fruitful Kickoff meeting with Data-Processing Team of CAS: will work intensely on two example projects for offline data processing of photon beam diagnostic data
Successful commissioning at 7.2 keV for XGMs at XTD2 and XTD9 (lower than usual SASE1 photon energy)

## Slide 8: CAS

## Slide 9: Detectors

## Slide 10: Advanced Electronics
Power glitches on Thursday morning:
  the PLCs and MircoTCA systems were not directly affected
  However, the received timing data from DESY was corrupted for a longer time
  As a result the PLC systems switched to simulated train ID mode and were switched back after the situation was stable again
Support to SQS: troubleshooting a non-responsive motor and assistance with encoder configuration

## Slide 11: ITDM
no issues related to the power cut
- sufficient redundancy is implemented
on Tuesday observed performance issues with GPFS storage
caused by malfunctioning InfiniBand switch
restart of the switch cured the problem