## Slide 1: SCS
Progress
TIM transfer measurement done
LIN: imager added
All imagers installed ->> wait for cables installation by actemium
permanent beamline ready for first beam next week
motors of KBS instantiated
Problems discussed already with involved groups

## Slide 2: SQS
Progress
Last vacuum connection between the instrument and the transfer beampipe has been done.
Installation of the target for the radiation test latest on Monday.
The vacuum in the KB mirror vessel is not good enough yet. Additional leak checks and RGA measurements have been performed but did not find a problem.
Motor tests on the laser in-coupling section have been done.
AQS chamber is now under vacuum.
The wavefront sensor was set-up (waiting for vacuum connection).
Last cables will be pulled until Tuesday noon. The chicanes will be closed in the afternoon.
The SQS hutches will be closed due to laser interlock tests from 16.00 on Monday.
Radiation safety tests on Thursday.

## Slide 3: MID
Progress

## Slide 4: HED
Progress
MEA marked drilling pattern for IA2 rails in EXP room A.12
MEA can spend a few hours next week to align optics hutch components
OPT hutch: We assembled PAM vacuum chamber (manipulators, ion pumps) and are ready to to local vacuum test of CRL, Attenuator, and PAM.
Cabling of Amplitude improved, TS to switch power back on (?)

## Slide 5: SASE2 XTD6

## Slide 6: Laser
SASE 3:
  Lasermet Interlock tests now to finish next week
  Electronics short circuted in control rack due to water leakage on top of E04! This needs to be prevented in the future.
  phase II installation otherwise complete.
  phase III will start with Amphos installation in August.
SASE 2:
  phase II finalization not clear yet (e. g. cooling water, under table installation, etc.).

## Slide 7: Vacuum
XTD6: issue with ion pump controller for M2 and MCP solved with VAC loop update yesterday
MID branch is under vacuum, gate valves are tested
HED installation is ongoing
Instrument support for MID, SQS, SPB...

## Slide 8: Optics

## Slide 9: Photon Diagnostics
MCPs
  MCP in XTD6: limit switches on all motors adjusted to allow for fail-safe operation at all now accessible motor positions in X and Y
  MCP in XTD6: unfortunately the HV-/Signal-test of BOS-MCP revealed that it might have a small crack and thus might be inoperational
  MCP in XTD6: MCP1 and MCP3 commissioned with UVlamp signal, ready for operation, MCP2 yet to be done
XGMs
  SCS microTCA inventoried (AE) and DOOCS server prepared (DESY)

## Slide 10: Advanced Electronics
Integration and tests in SASE2 tunnel
Developments on MC2 motor devices (e.g. homing)
Different activities at SQS and SCS (motor configuration, loop updates, testing)
Meeting with Martin Teichmann took place on Monday to clarify about communication issues
Invitation for meeting with HED/MID and different support groups for upcoming PLC and MicroTCA installation, testing and commissioning

## Slide 11: ITDM
SASE3 - configuration of filesystems ongoing
configuration of utility nodes
Activation of the network in SASE2 - work scheduled for Monday and Tuesday

## Slide 12: Detectors
2nd AGIPD MID
  AGIPD is expected be ready for transportation to XFEL end of August
  It has been communicated to the AGIPD consortium that we transport the detector to XFEL as soon as all modules have seen X-rays at DESY (requires detector to be present at DESY)
  Invited AGIPD project manager for a meeting with XFEL management to clarify the installation, integration, commissioning and calibration plans for the AGIPD detectors for XFEL
DSSC
  Testing of the DSSC vessel
    Access to the CFEL lab granted if we log people and activities. The problem of opening the door stays, but is probably minor
    Leak detector prepared for transport at CFEL, but need confirmation from CFEL people that they can be there when we plan to transport it
    Power supplies for lab MB proposal submitted, order for the chiller also submitted
    Three PPT feed through flanges tested and found leak tight, repairing procedure works
FastCCD
  Calibration measurements ongoing after vacuum problems with PulXar

## Slide 13: CAS

## Slide 14: SRP
July 26, 8am Commissioning of SASE 3 Radiation Hutches:
Machine runs single pulse
Steer the beam onto the SQS shutter (no focus).
At this time the SQS shutter can be opened.
People at SQS instrument have to check that the beam hit on the YAG screen in the hutch. If not it needs to be steered with tunnel mirrors (M2).
Enter the hutch and insert the target (manual translator)
The machine can go to 60 pulses. Start of radiation tests at SQS
Switch back to single pulse then steer the beam to the SCS branch (inserting M5).
Same procedure as before: Make sure that the YAG screen is hit properly, insert target (here remotely possible).
Switch to 60 pulses, start of radiation tests.

## Slide 15: SRP
The site inspection that took place in SASE2 today was sucessful. There are no points left hindering us to lift the PPE requirement for the area. This come into effect starting on Monday 23.07.2018
PPE is still required for craning (helmet and safety shoes), transport of heavy material and installation of equipment (minimum safety shoes and bump cap if necessary). In the workshop area the related rules are still in effect.
Sturdy shoes are a permanent requirement. High heels, sandals and and similar kind of shoes are forbidden due to ongoing work.
Due to the foreseen construction work in some fenced off and marked areas other PPE requirements can be in effect signposted at the barriers. (construction of SASE1 rackroom D.23, construction of cryo feedline SASE2, transport of laser system parts into A.23, ...)