## Slide 1: Photon Run Coordinator

## Slide 2: SPB/SFX
Activity report
- Optics hutch shutter unblocking and related interlock tests happening Fri morning (now)
- Loop 1 updated to contain coordinated motion / peltier cooling etc.
- vacuum is established in sample chamber via Ebara pump in pump-room
- vacuum in experiment hutch mirror tank at acceptable level for beam
- sample area re-cabling mostly completed
- vacuum issue at MKB_PSLIT (not properly closed hand valve) re-pumping with cart now, expect beam allowing vacuum level today (Friday)
- AGIPD work ongoing, now more stable, more pulses available. re-training required for those participating in first beam weeks

## Slide 3: SPB/SFX
To do’s:
AE involved (priority order)
- Pulse picker basic functionality is required for the mirror operations (must let through one pulse & one pulse only)
- Signal to one encoder is faulty MVE/Y_COARSE_DOWNSTRM_RIGHT, cable MKB_MVE.W4 (encoder itself is ok as can be powered from other cables)
- update for zoom/ focus of OHD_SCR (requires loop 2 update and re-wiring??)
- Some shutter related beam interlock questions not finalised (already in discussion)
- Some issues with Ion pumps in loop 4: could not move to ‘active’ from ‘passive’ (perhaps this is expected behaviour for not at setpoint?), now cannot get controller out of error. pressure value according to Karabo does not match controller reported value in either case
CAS involved (priority order)
- need middle layer device for handing of coupling / uncoupling and calculation of pitch angles etc. In development by G Giovanetti but has to be handed over to A Silenzi due to vacation. This is a critical device for mirror commissioning
- A GUI client update is required, an issue of missing scenes could have been avoided, update window available today (Friday)
- Basler cameras have to be re-started after re-cabling and require CAS assistance to communicate with the network/ Karabo

## Slide 4: FXE
VHS alignment for Br Kb, and vt also fits Ir La
LPD azimuthal integration efforts with CAS
Instrument alignment over weekend

## Slide 5: Laser
SASE 1:
100kHz tests stopped due to FE-damage, not allowing it to operate at this rep-rate
1.1MHz operation established.
1030nm beam delivered to SPB for upcoming user operation
1030nm beam delivery to FXE for upcoming user operation in progress (end of week)
NOPA setup for upcoming FXE user operation ongoing, some issues to be resolved

## Slide 6: Vacuum
SASE1 and SASE3 vacuum systems are ready to take beam,
All SAS1/3 tunnels closed
FXE Mono chamber opened last Friday
SQS transfer pipe connected to tunnel vacuum
SCS HSLIT shifted 6 mm to the right for radiation tests on Thursday, together with PRC (Manuel)
Problems with SASE3 absorber (flow control valve adjusted)

## Slide 7: Photon Diagnostics
General
photon commissioning shifts week 30 and 31 by XPD team.
XGMs
XTD2 and XTD10 XGMs are online XTD9 XGM will be online from Monday SASE3 reached > 4 mJ and the XTD10 XGM needed some adjustment Start of installation of XGM equipment in SCS

## Slide 8: CAS

## Slide 9: Detectors
1st AGIPD SPB
  Detector tests after the firmware upgrade (176 images per train)
    Detector performance w/o X-rays (no VETO) was tested and preliminary results did not show any issue related to the faster readout
    Operation scenarios for dark, X-ray, pulse capacitor have been prepared and data was taken for three configurations:
      No VETO (4.5Mhz operation)
      VETO (1MHz operation)
      Skipping unwanted pulses (1MHz and 0.5 MHz operation)
    First calibration data w/o X-rays (i.e. darks and pulse capacitor ) was taken for 176 memory cells  data analysis ongoing
  Stability of the system has been improved
    Readout FPGA firmware for one module Q2M3 was updated to fix the problem with ADC tuning
    Long term stability tests show that the interlock signal (i.e. MCLIVE) is stable and does not trip the detector power
  New uTCA crate (C&C system for AGIPD) was installed and configured  Karabo devices have to be created
  Work on operation manual ongoing  training for AGIPD operators will be scheduled soon (6-8.08)
  Plans for the next weeks
    Test external VETO as soon as new uTCA is integrated in Karabo
    Continue data analysis
    Prepare needed configurations and software for the next commission beam time
    Vacuum leak test with the leak detector as soon as the pump card is available

## Slide 10: Detectors
LPD FXE
  Plans for the next week:
    Testing the control system with Karabo 2.2.4
    Testing the new functionality in the calibration pipeline (i.e. radial integration)
    Testing the calibration pipeline with calibration catalog
    Take charge Injection data or all 510 cells

## Slide 11: Advanced Electronics
SASE3
Ion pump controller broken, support given
Pitch motors of the mirrors were wrong configured, now it is fixed.
Absorber in SASE3 VAC without pressure to operate, so therefore they were not able to open it.
Gas attenuator was updated
SASE1
SPB-CRLS motor lenses disconnected. Photo Run Coordinators were notified by Richard Bean, so they can check the MPS communication and beam mode interlocks.
Fxe MONO XTD9 - Different measures and checks were done and after talking with the manufacturer, the conclusion is that the motors are broken. Different vendors were contacted in order to get new offers as soon as possible.
SPB
  SPB01 – coordinated motion added
  SPB04 – Last part of coordinated motion pending to be added tomorrow 27.07 or Monday 30.07
  SPB02 – Technosoft pending to be tested and integrated.
  SPB05 – SHAFT encoder wrong configured, now working better but we still need more data to be able to configure it properly (Belongs to Z_STEPPER MOTOR)

## Slide 12: ITDM
new version of DAQ installed on SASE1 and SASE3 tunnels, instruments will follow
rearrangements of raw and calibrated data
calibrated data from 201701 copied to dCache
all proposals from 201701/201830 are closed
consistency checks are ongoing
removal of raw and calibrated data from offline GPFS will take place on 6.08.2018
announcement was posted to computing@xfel.eu mailing list
PIs/Main proposers of the affected proposals informed by direct email from User Office