## Slide 1: T-mapping, B-measurements and Vertical Diagnostics at Cornell
Ryan Porter
Cornell University
Supported by:
U.S. DOE award DE-SC0008431.

## Slide 2: Outline
Vertical test diagnostic
  Setup
  Analysis
Temperature mapping
Flux Trapping Measurements
  Experimental setup
  Comparison with other labs
  Electromagnetic simulations

## Slide 3: Experimental setup
Cernox Temperature Sensors
Flux Gates
Helmholtz Coil
3 Cernox sensor
Bartington Flux Gates (Mag-01H type F)
  Historically:
    Always one of top iris
    Sometimes another on equator
Helmholtz coil

## Slide 4: Phase Locked-Loop
B
F
Diagram take from: D. L. Hall, “New Insights into the Limitations on the Efficiency and Achievable Gradients in Nb3Sn SRF Cavities,” Ph.D. Dissertation, Cornell University, 2017

## Slide 5: Measurement Techniques
Quality factor
  Use decay curve method
    Use decay constant near peak field to calculate Q
  Analyze entire decay curve to find Q from peak Eacc to low field
    Allows to measure Q vs E at many temperatures during cooldown
    Confirm by measuring Q vs E using only peak field at some temperatures
f vs T
  Measure frequency of cavity vs temperature
    Remove effects from pressure
  Acquire change in penetration depth by fitting

## Slide 6: Coupling Calculation
Use average of two coupling calculation
Measure Pf and Pr with power on then Pe (emitted power from cavity) immediately after turning off power to the cavity.
  Beta is calculated from Pe and Pf:
  Beta is calculated from ratio of Pf and Pr
    Where + is chosen if Pf < Pe and – is chosen if Pf > Pe
Only accept measurement if the two couplings are close in value

## Slide 7: Temperature Mapping at Cornell

## Slide 8: Temperature mapping
T-Map board
646 sensors on cavity + 4 LHe bath sensors
~1.7 cm between sensors on equator and flats of cavity
16 bit ADC at 24 kHz
Multiplexed
Typically average 1000 measurements for each resistor in T-map
Takes several seconds to read entire board
Can also measure a single sensor at 24 kHz

## Slide 9: Temperature mapping
T-Map sensor
T-Map sensor head
Niobium surface
Thermal paste
Pogo stick
Carbon resistor
Allen-Bradley carbon resistors
  Bad resistor = good thermometer
Resistors embedded in stycast to shield from LHe bath
Shielding on resistor removed to allow direct contact to cavity wall
Springed “pogo” sticks keep sensor in contact with cavity surface
Silicon grease applied to sensor as a thermal paste
Connector leads

## Slide 10: Calibrating T-Map
Measure resistance by applying constant current source and measuring voltage drop
Allen-Bradley resistors ~30 Ω/mK
  Need calibrating
Calibrate while cooling below 4.2 K using Cernox sensors
  Fit a 4th order polynomial to T-1 vs ln(R)
Can get 1 mK resolution (after averaging)

## Slide 11: Example Temperature Map
Temperature map of a Nb cavity at Cornell
ΔT = TRF-on - TRF-off
Use bath resistors to account for LHe bath T drift

## Slide 12: Flux Trapping Measurements at Cornell

## Slide 13: Experimental setup
Cernox Temperature Sensors
Flux Gate on top iris
Helmholtz Coil
Cooldown in applied field, then turn off Helmholtz coil
Btrapped = Bleft - Bamb

## Slide 14: Three Methods
4
Method adopted at JLAB[1]
Method adopted at FNAL[2]
Method adopted at Cornell[3]
[1]. S. Huang, Takayuki Kubo, and R.L. Geng, Phys, Rev. ST Accel. Beams 19, 082001(2016).
[2]. M. Martinello et al., in Proceedings of SRF2015, Whistler, BC, Canada, MOPB015.
[3]. D. Gonnella, J. Kaufman, and M. Liepe, J. Appl. Phys119, 073904 (2016).
Method J
Method F
Method C
Numerically Calculated Ratio
Modeling perfect dia-gmagnetism
RPD = 1.7 for 1-cell TESLA end cell shape
Varies with cavity geometry
Shichun Huang(IMP,CAS), Rongli Geng(JLAB), “Can Present Techniques for Cavity Flux Expulsion Efficacy Measurements be Unified?”, TTC 2018 CERN, Nov 2018

## Slide 15: Flux Trapping Experiment on STE1-2
Was preparing to test STE1-2
  2.6 GHz Nb3Sn cavity
    Will present results on Thursday!
Opportunity to compare flux trapping techniques!
Compare techniques all in one test with top iris, bottom iris, and equator flux gates
Cornell:
  Cooldown, measure top iris
Jlab:
  Cooldown, measure equator
  Cooldown in “zero” (ambient field) then turn on Helmholtz coil to same current
Fermi Lab*:
  Cooldown, measure equator
  *Calculate/simulate new RPD number for specific geometry
Note: cooldown rate 1 K/10 min

## Slide 16: Flux Trapping Experiment on STE1-2
STE1-2 looks like it is near 100% trapping
  For 106 mG applied field change in field through transition/transitions is 1 mG or less.
After cooldown with Helmholtz coil off flux top iris fluxgate read 46.3 mG
  Much less than 106 mG!
Historic Cornell method under estimates trapped flux
“Real” | JLab | Fermi | Cornell
106 mG | 106 mG | 106 mG | 46.3 mG

## Slide 17: Simulations
Simulated STE1-2 flux trapping/expulsion
  Model in CST Studio with very small µr (0.0000001) (near perfect diamagnet)
  Complete Expulsion of 106 mG
  106 mG trapped, no Helmholtz coil
    Assumption Field trapped perfectly vertical
Average B-field over active region of flux gate
100% Expelling (cooled in 100 mG)
y-component B-field
100 mG trapped
y-component B-field

## Slide 18: Simulation Results
Simulation results for 106 mG applied/trapped
  Within 15% for all measurements!
  ‘Tabs’ are important in simulation (for 2.6 GHz)!

## Slide 19: Equator is bad… Why!?
Simulation not entirely getting correct fields inside walls. Field switches from 100 mG at halfway point of wall (1.25 mm/2.5 mm)
  High density mesh or smaller µr does not help
  Problem with program for this application?
Magnetic field incorrect through half of wall

## Slide 20: LTE – 1.3 GHz Tesla Style Cavities
Have conducted simulations of LTE (1.3 GHz) flux trapping
  The cavity simulations find top iris fluxgate measures 58.0% of ‘actual’ trapped flux
Previous data could be corrected using simulation results
Need to confirm flux trapping simulations with experiments

## Slide 21: Flux Trapping Sim Conclusions
Cornell’s historic iris measurements trapped flux have underestimated trapped magnetic fields
  Cornell measurements have overestimated flux trapping sensitivity
Some improvement needed in flux trapping simulations
  B-field changing in walls
For STE (2.6 GHz) cavity top iris measures only 43.7% of ‘actual’ trapped flux
For LTE (1.3 GHz) cavity simulations find top iris fluxgate measures 58.0% of ‘actual’ trapped flux
  Need to do 1.3 GHz flux trapping experiments to confirm simulations

## Slide 22: Further Questions
Flux trapping measurements assume trapped flux is oriented with the applied field. If trapped flux moves during partial expulsion, how much does this impact measurements? T-maps imply that the flux distribution is not uniform.
How does the orientation of trapped flux w.r.t. the RF surface impact losses? And sensitivity measurements?

## Slide 23: Back up

## Slide 24: Continuous Q vs E
For continuous Q vs E the entire decay curve is used
The conventional decay curve Q vs E point is found
This first point is used to determine Qe
This allows a continuous Q0 to be calculated as:
Eacc as a function of time is: