## Slide 1: Joint Operation & Readiness meeting
Report Operation / Photon Run Coordinator
Report Dispatch
Instruments
  FXE
  SPB/SFX
  SCS
  SQS*
  MID
  HED*
Beam transport
  Task force XTD6 installation*
  Vacuum*
  X-ray optics
  Photon diagnostics
Optical lasers*
Detectors*
Electronics
  AE*
  EETF
CAS
ITDM
SRP
Technical services
AoB

## Slide 2: PRC

## Slide 3: Report from Dispatch / Various

## Slide 4: SQS
Continuation of commissioning work WITHOUT beam

## Slide 5: HED
additional 2 racks on AC platform (A27) approved by safety, Thursday (yesterday) there was a first kick off meeting with PI
misson loop-6 CRDs to be ready in 1 week from now, need to be included in next cable order for day-1 experiments in May 2019.
DESY started drilling of floor for rail system and covers in A.12
Interlock TUEV for HED OPT (end-Sept 2018) had to becanceled, due to misinformation: We miss the link tube (MID-exp to HED-opt), which we were told by SRP is not needed, but DESY made a walk-though and requsted it. We need to clarify this.
OPT: finalized z-adjustments of components in OPT hutch, ready to order connection pipes (again, we were told by SRP that these were not time critical for TUEV tests).
Regulation of AC devices with heat load ready, VOB: Klima und Waser, Elektrik passed. But MSR (Messen Steuern Regeln) failed, must be repeated. Finish missing works by 21 September 2018

## Slide 6: XTD6 Task Force
See PDF on Indico

## Slide 7: VAC
SA3:
interlocks for Gas attenuator and Mirror tested successfully on Tuesday
probably another access to XTD10 for modifications on frontend absorber
SA2:
missing vacuum parts for HED branch are delivered, installation in progress
frontend-section for HED was closed and is under vacuum
preparation for installation of Mirror M3 and HMONO (HED)

## Slide 8: Optical Lasers
SASE 1:
PP-laser is in operation
SASE 3 and 2:
We are lacking support from CAS/AE/ITDM. The following installation critical things need to come asap to avoid delays:
ticket #24879 - Request for Control P.C.s (fully installed Karabo....) for Pump-Probe Laser Galvo in room E.04 and for USB devices in room A.04
ticket #24833 - Fwd: Ethernet Switches and Patch Cable for SASE2 and 3 Laser Hutches and Laser Lab
PLC hardware and firmware installations, extension crate firmware in SASE2 and 3 Laser Hutches and Laser Lab
Timing fiber

## Slide 9: Detectors – DET
1st AGIPD SPB
  Operation – User proposal p2120:
    Configuration with 128 memory cells, 120 X-ray pulses per train @ 1Mhz were collected
    Collected data: 123 TB of raw data
    Issues:
      Configuration of the detector was not always properly applied  the detector operators were made aware of this problem and temporary solution (before the problem will be fixed on the firmware level) was provided
  Preparation for commissioning beam time ongoing
LPD at FXE
  Operation
    Successfully finished last beam time
  Issues
    Installed Karabo device to restart calibration pipeline in case of problems -> to be tested

## Slide 10: Detectors – DET
2nd AGIPD MID
  Detector is pumped P=2x10^-6 mbar @ Temp=-10degC on cooling blocks
  It was discovered that the radiators used in the system are NOT suitable for pressure of 3 bars.  the new radiators were installed and leak test was performed
  Interlocks are operational: Interlock signals from detector(i.e. MC-LIVE and T_LTCC_OK) was tested.
  LV power configuration was adjusted for the MID AGIPD and the detector was powered up (up to ASICS<=> LV)
  Control of AGIPD via “old” Karabo (2.1.1) was tested successfully
  Karabo was upgraded to Karabo 2.2.4 and successfully tested
  Issues
    MFPGA board of wing 2 does not have sense wire connection
    MC_LIVE signal trips interlocks (the same as it was at SPB before the firmware on micro Controller was upgraded)
    Two ASIC power channels show some issues --> problem under investigation. At the moment, I power these channels w/o sense wire feedback, so they are underpowered.

## Slide 11: Detectors – DET
FastCCD
  Testing and bug fixing of new control software ongoing
  In vacuum cables are on the critical path. LBNL informed, will take all measures to speed up delivery to XFEL until end of next week.
DSSC
  Testing of the DSSC vessel
    Quadrant mounting tool checked, several flaws identified. Fixing of the design has started. Company identified which could deliver the missing parts in three weeks. Most of the effort is now put in minimizing the delays.
    Revised commissioning document (with remaining tasks to be performed at CFEL) sent to CFEL personnel, almost finalised.
    Transport of the vessel from CFEL/DESY to HERA South planned on Friday next week.
  First quadrant ready at DESY, could be installed middle of September. Second quadrant will be assembled by the end of the week.
  PLC systems
    All (but for SIB needing modification) in house and ready. Installation will start next week.
  PETRA proposal for system calibration submitted

## Slide 12: AE
SASE2 Tunnel:
Crates finished and located in the racks for HED loop. Benoit has provided eplan export and PLC will be up beginning next week.
MID BMPIs tested, mechanical issues detected and therefore BMPI-2 Screen should be recommissioned after checks are done.
SASE3 Tunnel:
Interlocks update for Vacuum and EPS to MPS interlock. Has to be tested by Harald Sinn following week.
Issues with parameters of motors lost, working in a solution to restore them after PLC reboot.
SCS:
Smaract stages working properly but still one crate with smaract cards missing.
One smaract card has problems, but could be due to cable issue.

## Slide 13: AE
SQS:
Loop 6 is running but still some motors disconnected and Technosofts motors need to be downgraded.
SPB:
Smaract stages reconfigured and working.
SPB4: Ion pump controller working properly for first channel and change request expected in the crate to be able to read channel 3.
Injection movement with one faulty limit switch that should be replaced.
MID:
Working in PLC loop 1.
MPS interlock definitions:
Work flow should be properly set between XFEL staff before talking to D3 people (meeting will take place).