## Slide 1: Joint Operation & Readiness meeting
Report Operation / Photon Run Coordinator *
Report Dispatch
Instruments
  FXE
  SPB/SFX
  SCS*
  SQS
  MID
  HED*
Beam transport
  Task force XTD6 installation*
  Vacuum*
  X-ray optics
  Photon diagnostics
Optical lasers*
Detectors*
Electronics
  AE*
  EETF
CAS*
ITDM
SRP
Technical services
AoB

## Slide 2: PRC Report

## Slide 3: SCS
Activities last week:
Loop 1 Interlock installation is ongoing
XGM cabling is ongoing
Installation of sample changer is ongoing
bake out of the KB was stopped; pressure is at 3 E-8 mbar
transfer pipe were installed and pumped; pressure is at 1 E-9 mbar
TIM was connected to rough vacuum, ready to install the MCP
8 digitizers are installed from AE
HV supply for MCP was installed from AE
Problems:
chiller situation for DSSC unclear; same problem like with the julabo
chiller for AGIP?

## Slide 4: HED
XTD6: Grouting of hrMONO plate successful, delivery of hrMONO on Monday.
EXP: drilling for rails continued, almost finished. Mount first cover parts next week.
TUEV test of OPT and EXP hutch postponed by HED to early Feb 2019. However DESY D3 is not available entire Jan-Feb 2019, so we may have to complete the OPT hutch in December 2018. (we need to commission OPT in Feb 2019)
Temporary solution for liquid N2 in A.23 could not be guaranteed by TS by 1 Oct 2018. Therefore we canceled current installation schedule with Amplitude. We will now install the final fixed lN2 supply line (Dräger, Cryoterm - should have been ready Sept-1). This shifts the SAT of the Amplitude laser from Nov 2018 to Jan-Feb 2019 and also we will NOT include the laser for users in call 4.

## Slide 5: XTD6 Task Force
See PDF in Indico

## Slide 6: VAC
SASE1 and XTD1 taking beam,
All SASE1/3, XTD1 tunnels closed
Preparation for "Mini-Shutdown” next week:
- FXE Monos
- XPD work on HIREX and Imager in XTD9
SASE2 vacuum system was completely closed for short time
HED HMONO connection to vacuum system next week
M3 installation today
Exchange of terminals/loop update in XTD1 to fix “SA2_VALVES_OPEN” signal issue planned for Wednesday
Tunnels will be opened:
XTD1,2,4,10: Monday to Wednesday
XTD6,9: Monday to Friday, or even longer…

## Slide 7: Optical Lasers
SASE 1:
PP-laser operation as planned
SASE 3:
Seeder is running
Front-End is being powered up and tested
under-table water installation finished
PA seed beam line is being prepared
humidity tests were done
SASE 2:
under table cabling nearly finished

## Slide 8: Detectors – DET
1st AGIPD SPB
  Operation – User proposal p2119 (13-17.09):
    Configuration with 4 memory cells, 1 X-ray pulses per train
    Collected data: 2.4 TB
  Preparation for shutdown (24.09-09.10):
    Task list agreed with SPB and AGIPD team
  Issues:
    Water in the Si-oil cooling circuit did not allow to cool down AGIPD detector  running the system @ 45 degC for ~15 hours solved (at least temporarily) the issue
    The detector baseline shift due to high intensity signal  work on the corrections ongoing, not yet implemented in the calibration pipeline
  Operation – User proposal p2114 (20-23.09):
    Configuration with 128 memory cells, 120 X-ray pulse
LPD at FXE
  No issues to report, preparing for beam time

## Slide 9: Detectors – DET
2nd AGIPD MID
  AGIPD data was recorded to the hdf5 for 8 modules at the same time
  DAQ was upgraded for recording data at the same time from all 16 modules  tests ongoing
  Integration of Clock& Control Karabo devices for AGIPD successfully tested
DSSC
  DSSC vessel successfully transferred to HERA South hall
    Vacuum system was installed in HERA South, the vessel is pumped since Monday and the pressure is now 1.4 x 10-6 mbar
    PLC system (motion and vacuum) installation ongoing, PLCs to go online on Thursday (today). Firmware installation not yet done, software installation by CAS started.
    Leaky lateral flanges under rework.
  Power supply control in Karabo
    Tables produced by CAS, new device tested, training in preparation for person to test the system.
    Test at XFEL will start after PLC installation.
  First quadrant ready at DESY, could in principle be installed in the already. Second quadrant in preparation.
    Quadrant mounting tool order went out this week. Company promised three weeks production time.

## Slide 10: Detectors – DET
FastCCD
  Testing and bug fixing of new control software ongoing
  In vacuum cables
    Electrically tested in vacuum cables and mounted them
    CCD performance tests with long cables are presently ongoing
  Tested mounting of the in vacuum filter holder/frame
pnCCD 1 Mpix
  Preparing performance tests with FastADC with support from AE and CAS

## Slide 11: AE - 1
PLC
SQS
MPS interlock definition being checked.
Loop 1- Problem with one ionpump and PLC has to be redone to fix it.
Loop 3 – Ionpumps channels wrong declared, eplan was corrected and we will generate the project again.
Loop 4 – some problem in wiring of analog sensors that will be checked by SQS, then we can test it.
Loop 5 – project being prepare with modified crates and new motor parameters.
SCS
Loop 1 running with interlocks – has to be tested by SCS.
Loop 7 and 8 running – they should be tested by SCS.
SASE3 PPL
PLC ready, following week we can start to generate the project and try to run it.

## Slide 12: AE - 2
SASE2
HED CRL still in commissioning face
HED PPU pre-commissioning on Thursday, project will be prepared by next week.
HED MONO – Chris will request to test the in-vacuum motors before chamber is closed
MID. - Leds for BMPI will be integrated next week
MPS changes in Vacuum and EPS scheduled for next week.
MID
Loop 1 being tested and commissioned
Loop 4 running, needs testing
List of priorities received from MID

## Slide 13: AE - 3
MicroTCA
MID
MicroTCA related cables are being routed in the rack room by the group, once this is done we can start configuring the crate
SQS
Preparing necessary crates for X-Ray/Laser/Rack room
SCS
Installation of the new ADQ7 digitizers in the rack room
Laser
Ongoing migration of Laser3 MicroTCA Crate (Timing done, digitizers next)
Ongoing installation of MicroTCA Crate for the Laser Lab 0.01

## Slide 14: CAS
See attached presentation