## Slide 1: Joint Operation & Readiness meeting
Report Operation / Photon Run Coordinator
Report Dispatch*
Instruments
  FXE
  SPB/SFX
  SCS
  SQS
  MID*
  HED*
Beam transport
  Task force XTD6 installation*
  Vacuum*
  X-ray optics
  Photon diagnostics
Optical lasers*
Detectors*
Electronics
  AE*
  EETF
CAS
ITDM
SRP
Technical services
AoB

## Slide 2: Report from Dispatch meeting
FXE imagers cannot be seen by the people in the BKR and that's a problem to efficiently communicating with them about e.g. beam stability issues. According to CAS a Karabo client is available for that purpose. The issue should be addressed by the PRC with the help of CAS.  Tbd.

## Slide 3: HED
High resolution Mono from FMB in XTD6 assembled, aligned. SAT ongoing.
HIBEF Rail system: cover construction aligned.
Draft of steel construction for the emergency exit A24 available.

## Slide 4
MID
Progress week 39 (Sept. 24-28):
TÜV test of MID OPT successful
Repairs/replacement of CRL-2 parts (XTD6)
Installation of AGIPD cooling pipes in drag chain (MID EXP)
Karabo/firmware tests (tech commissioning) of MID OPT almost complete
Karabo tests of imagers (XTD6) in progress
Other:
Installation of AGIPD cables (cable mountain) in drag chain urgently needed
Preparing to install: laser table, mirror chamber, DES, and DPS

## Slide 5
MID
Rough time plan:
SASE-2 lasing, XTD1 and XTD6 commissioning with beam, week 41-42
1st beam at MID and radiation tests of OPT week 43, installation of DES in EXP
AGIPD cables installed in drag chain must finish week 44
Installation of laser table and differential pump in EXP week 45
Installation of mirror vessel in EXP week 46
Installation of AGIPD latest during week 46
TÜV test MID EXP week 47
1st beam in MID EXP week 49
Open questions:
What is needed for the TÜV test of EXP week 47? (ongoing discussions...)
Is the AGIPD delivery schedule realistic?

## Slide 6: XTD6 Task Force
See PDF in Indico

## Slide 7: VAC
FXE Monos work ongoing, goal is to exchange linear stage of Mono1 by Thursday next week
XPD work on HIREX done
XPD imager will be pumped and leak tested together with Monos next week
HED HMONO connection to vacuum system today
M3 chamber closed

## Slide 8: Optical Lasers
SASE 1:
successful user operation to date
FE service performed
recovering from power outage
SASE 3:
Installation progressing
SASE 2:
yet another flooding from above. This has to stop, or we will face severe damage, lack of progress and interruptions of future user operation.

## Slide 9: Detectors – DET
1st AGIPD SPB
  Problem with in-vacuum cooling (i.e. ice in the Si oil) was fixed by running the Julabo for 15 hours at +45 °C
  Potential solutions using closed loop cooling concept were discussed with beam line and Huber representative
  User Operation
    Problems with application of the configuration to the detector
  Shutdown (24.09-09.10):
    SPB was asked to perform vacuum leak test this week.
    electronics in wing one (i.e. ADC boards for Q1M2 and Q1M4) were removed from the system and repaired. The system was tested after this interventions.
    SPB works on installation of new cooling connection for wing 1 electronics (it has to be done before the system is moved to the downstream position)
    Task list agreed with SPB and AGIPD team
LPD at FXE
  No issues to report

## Slide 10: Detectors – DET
2nd AGIPD MID
  Possibility of recording all 16 modules at the same time was sucessfully tested
  Raw online preview was installed but not yet tested
  First ramp up of HV was done this morning
DSSC
  DSSC at HERA South hall
    PLC crate installation being finalised, two missing power supplies to be put in. At the moment, vacuum and motion systems PLC are available, f/w and s/w installation to follow.
    Preparation work for connection to chiller started
    Investigation of connection to the motion system, contrary to what communicated the feedthrough flanges and the encoder cables have the same gender, need to remove the connector and install one with the right gender (next week probably).
    Adapter cables for motion system interface ordered and in preparation.
    Leaky lateral flanges still under rework.
  Power supply control in Karabo
    Device exists and is tested, first tests will start on KW 41.

## Slide 11: Detectors – DET
  First quadrant ready at DESY. Second quadrant in preparation.
    Quadrant mounting tool under preparation at the company.
    Next week work on quadrant assembly will restart at DESY
FastCCD
  Testing and bug fixing of new control software ongoing
  Discovered issue related to the mounting of the cryo cooler, requires mechanical modification of vacuum cross
1 Mpix pnCCD
  Nothing to report

## Slide 12: AE - 1
PLCs: SQS: Working on loops 7/10/14
SCS: Working on different loops, upcoming loop 8.
SASE3 Tunnel: --Vacuum updated. --Working on Master loop to forward shutter signals to instruments.
SPB/FXE: Whole Thursday checking PLCs and bringing up them after power shut down. We will need more communication or better understanding of the installation to try to be more efficient in these maintenance tasks.

## Slide 13: AE - 2
MID: Also lost of power in PLCs racks.
Framework release 1.31.0: --SD_LED is now included and will be tested tomorrow in SASE2 MID loop. --Homing procedure for SD_MC2beckhoff motors is included in the new release and will be tested also in this MID loop. There is a change in the way the homing will work, so documentation and explanation will be provide to the users.
MicroTCA:
MID - MicroTCA Crate completely configure (timing synchronization perform, all systems working)
SCS - Installation of the SIS8300 Digitizer w/ Stretcher RTM - MPS fiber placed and connected. Functionality confirm by DESY