## Slide 1: Joint Operation & Readiness meeting
Report Operation / Photon Run Coordinator*
Report Dispatch
Instruments
  FXE
  SPB/SFX
  SCS
  SQS
  MID
  HED
Beam transport
  Vacuum
  X-ray optics
  Photon diagnostics*
Optical lasers
Detectors
Electronics
  AE
  EETF
CAS
ITDM
SRP
Technical services
AoB

## Slide 2: Report from PRC
General Operation
4am tuesday the accelerator operation was stopped to repair a cold box motor in the cryo-system. This shortened the FXE user run by 2h and made some rearrangement and cancellation of shifts necessary.
FEL beam was back in all three SASEs early Wednesday afternoon, restart of photon system activities
Beam delivery to all 3 SASE beamlines:
SASE1: preparations for user runs in FXE and SPB; users starting friday day shift in FXE
SASE2: commissioning of XTD6 (from BKR) :
MID branch advanced, HED branch ongoing this week.
All imagers and MCP set to correct timing,much scene development.
Much work completed on mirrors and MID-imagers. M3 coupled motion this weekend
Started commissioning of MCP, XGM, imagers, PPU, ...
Plans for Kmono measurments
SASE3: commissioning in SCS and SQS
need to get SCS and SQS shift summaries copied into the operation ELOG ("SASE1-2018") !

## Slide 3: Report from PRC-2
Issue: Interference of Pulse patterns or number of bunches in the SA1-SA3 north branch, e.g. during interventions like screen insertion or PoD.
Example: FXE receives 100 bunches, concurrently SQS gets 5 bunches. FXE drives in an imager which - thanks to EPS - changes the mode of the machine to 2-bunches for FXE. Since in this mode the machine can deliver max. 2 bunches to the north branch, then SASE3 in that moment doesn't get any lasing bunches anymore AT ALL. Even worse: when the imager will be retracted sometime later, unexpectedly and suddenly SA3 receives 5 bunches. The unpreparedness can even cause damage in SA3, if the time of no-beam is substantially long.
Short-term remedy (communication!!): FXE requests from BKR 1 bunch for SA1 subtrain, then BKR can set also 1 bunch for SA3 subtrain, and thus both can work happily with imagers etc. for alignment. Longer-term remedy: instead of forcing something into the beam that by EPS-protection throws the machine into a different-bunch mode, rather the machine bunch pattern must first be changed intelligently (either by calling BKR or by a script as used for PoD at FXE). The EPS-protection should only used for its primary purpose: to protect equipment against damage.

## Slide 4: Report from PRC-3
upon several PLC loop updates during last two months, slits in SASE3 was tested first time last week.
(i) Exit Slit SCS - speeds of two blades were different and much higher (0.4 and 0.1 instead of 0.05), although was reported by AE and CAS colleagues that everything is OK. In this way, the middle layer device which is capable to move gap position, means move blades in one direction, can bring the blades into collision. How to improve: (a) check for speed in middle layer device; (b) improve procedure of recovering parameters from PLC and Karabo side.
(ii) Vertical Slit has lost both motor positions and encoder positions (encoders are incremental). When trying to move, one of two blades was moving in one direction, in the beam. Changing parameters, namely operation mode to 1 helped (but I did not dare to continue checking on Sunday). Still, calibration of width is completely lost, can be done with beam only by diffraction, will take sometime. How to improve: (a) to change hardware to absolute encoders; (b) I have already reported to Patrick that operation mode 2 most probably leads to loosing of positions upon loop update - can be checked on PLC side

## Slide 5: HED
A23 high energy laser beam transport installations progressing
A12 high energy laser beam transport vacuum tube down the wall installed
Opt hutch: working to close vacuum in first half for TÜV test in Nov.

## Slide 6: XPD
Commissioning of Diagnostics in XTD6: overall good progress, but still much parameter setting and scene development necessary and ongoing
Beam delivery in all SASEs
XGMs:
Online: XTD1, XTD2, XTD6, XTD9, XTD10
On Saturday first light at XTD6 XGM (HED branch) ! - adjust timing of fast signals. One cabling issue found and repaired (together with DESY colleagues)
SCS-XGM: - Successful HV tests at the SCS XGM - Randomly disappearing PCIe devices MTCA: exchanged MCH (spare from AE), start long term test.
Imagers
XTD6 imagers beam commissioning started - all triggers adjusted
some few MID-imagers still have DAQ-issues (files and entries created but no images recorded)
MCP
XTD6 MCP beam commissioning started - all triggers adjusted
thursday: first time signals observed on all MCP detectors including visual BOS-MCP
some DAQ issues identified and resolved