## Slide 1: Joint Operation & Readiness meeting
Report Operation / Photon Run Coordinator
Report Dispatch
Instruments
  FXE
  SPB/SFX
  SCS
  SQS
  MID
  HED
Beam transport
  Vacuum
  X-ray optics
  Photon diagnostics
Optical lasers
Detectors
Electronics
  AE
  EETF
CAS
ITDM
SRP
Technical services
AoB

## Slide 2: XPD
General XPD:
• Strong diagnostics support to the Machine Study Beamtime by E.Shneydmiller :
  o Five people from XPD at BKR all day during the holiday 31.10.2018
o Several issues like a hardware malfunction of the SASE3 hutches personnel interlock --> EPS-effect on gas attenuator, impeding most diagnostics
o Harmonic Lasing was unfortunately not successful
o Reverse Tapering worked !! (detection with SR imager) :
"The first 12 undulators have inverse tapering 0.2% per cell. No SASE. If we close additional 2 undulators and tune them to resonance, than we have 300 uJ at 1.5 nm.“ More details, see ttfinfo-logbook: https://ttfinfo.desy.de/XFELelog/show.jsp?dir=/2018/44/31.10_a&pos=2018-10-31T22:59:59
• Teamcenter training (2-days) done for 4 people of XPD
XGMs
• Online: XTD1, XTD2, XTD6, XTD10, SCS.
• Offline: XTD9 (due to tunnel access)
• SCS: found fast and slow signal. Will need Xenon when Monochromator is used because of strongly reduced intensity.
Imagers
• Study following effect: Photonic Science cameras stop sometimes spontaneously at the same time in all three SASEs (mains? trigger? network?)
PES
• Design Reviews for Hard X-ray PES : Nov 05 10:00 – 12:30 (E2.042) and Nov 14 14:00 - 16:30
HIREX:
• Successful repeated application for the users during the last FXE user beamtime last weekend
Back

## Slide 3: Optical lasers
SASE 1:
- user run successfully finished
- PP-laser will be in maintenance mode and not available till before the commissioning run starting on 5th Dec.
SASE 3:
- Lasermet interlock commissioning in E04 ended, shutters now addressable and DACHS access works.
- incremental progress in PP-laser installation (Amphos and Front-End installation and commissioning)
SASE 2:
- regulated back cooling water should become available on 5th Nov.
- Witt / DESY interlock testing should start in same week
- otherwise incremental progress (cabling)
Back

## Slide 4: Detectors – DET
1st AGIPD SPB
  26-29.10: User beamtime (p2100)
    Configuration with 128 memory cells, 120 X-ray pulses per train
    Collected raw data: 112 TB
  Preparation for movement of the detector to upstream position  detector hole adjustment will be done on Friday
  Issues:
    No major issues
LPD at FXE
  No issues to report

## Slide 5: Detectors – DET
2nd AGIPD MID
  Problem with C&C signal for MFPGA board of wing 1 solved
  Installation of missing sense wire connection of MFPGA H2 done
  Firmware fine delays tuning will be done today afternoon
  Preparation for moving the detector from HERA south to XHQ ongoing:
    Preparation of the system for transportation 05-06.01
    Transportation of the detector from HERA South to Schenefeld 07.11
    Work on installation and commissioning schedule of the detector at MID instrument ongoing  try to clarify the status of MID infrastructure needed for detector operation
  Issues:
    Readiness of MID infrastructure for the detector operation  direct impact on the AGIPD commissioning schedule:
      PLC Loop 11 preparation is in progress, eplan export will be done in cw46  firmware, Karabo interfaces and ILC low level components tests will follow not ready before cw47
      Cooling: cooling circuit for Si-oil needs to be prepared. It is not expected to be ready before cw48.
      System cannot be powered w/o cooling and basic functionality of loop 11
      Non-AGIPD related work in MID hutch/rack room and its impact on the detector commissioning schedule has to be clarified
    Transportation of detector to XHQ  due to the construction work at Bahrenfeld, the road to HERA South is not accessible for LKW. It has to be clarified with DESY.
    Reworked quadrant mounting tool arrived, rails still missing, this prevents the testing

## Slide 6: Detectors – DET
DSSC
  Integration in HERA South
    PLC crate installation finalized.
    In-vacuum wiring of the motion system finalized
    Motion system being tested with manual control. Dummy quadrants will be installed on Friday or Monday.
    Reworked quadrant mounting tool arrived, partially tested, some parts needed reworking, partially done in house, the rest will be done outside.
  Power supply control in Karabo
    Under test, some issues identified
  First three quadrants are ready at DESY
  DSSC to be mounted and to some extent tested by the end of the year
    Tight schedule needs careful planning  done until the end of November
FastCCD
  Chiller modification to use DI water and/or interlock could delay commissioning
  Installation of camera on beam-line started today
  Electronics installation moved to November 6
  Commissioning at beam line to start after installation
Back

## Slide 7: Detectors – DET
JUNGFRAU
  Off-line calibration for Module #039 (installed at FXE) in place
    Offset subtraction
    Gain correction with PSI gain maps
  Completing analysis of SAT data
  Starting procurement of spare parts to repair 1M assembly