## Slide 1: Joint Operation & Readiness meeting
Report Operation / Photon Run Coordinator
Report Dispatch/PSPO
Instruments
  FXE
  SPB/SFX
  SCS
  SQS
  MID
  HED
Beam transport
  Vacuum
  X-ray optics
  Photon diagnostics
Optical lasers
Detectors
Electronics
  AE
  EETF
CAS
ITDM
SRP
Technical services
AoB

## Slide 2: Report from PRC
https://ttfinfo.desy.de/XFELelog/show.jsp?dir=/2018/46&pos=2018-11-13T17:31:54
https://ttfinfo.desy.de/XFELelog/show.jsp?dir=/doc/Schedule&pos=2018-11-12T09:51:45
Back

## Slide 3: Dispatch meeting report / PSPO
PSPO has now provided overviews regarding the winter shutdown activities @ XFEL (XHEXP1 and XTDs). All documents should be in this folder on Alfresco:
  https://docs.xfel.eu/share/page/site/PSC/documentlibrary#filter=path%7C%2FPLANNING%2FSHUTDOWN%2FWintershutdown%25202018-19%7C&page=1
You can find PDFs with activities filtered by location (individual XTDs or SASE-areas in the XHEXP1) in the folder "By Location", and PDFs filtered by requested support groups for AE, CAS, ITDM and VAC in "By Support Group"
Whenever an activity has meaningful start & end-dates, we displayed it in the Gantt-chart-style plot on the right bottom.
  Green: Activity should not affect other works,
  Yellow: Could disturb / affect some people, have a closer look,
  Red: Activity will affect an important central resource or service
We tried to have the major works, in particular red activities, in the maintenance period, such that commissioning from middle of January on should not be affected.
We ask you to check for potential collisions / conflicts of interest and give us feedback!
Back

## Slide 4: HED
Opened XTD6: Pulled cables and connected basler cameras to all 3 monochromator vessels
Cabling of OPT started by Fa. Schulz.
OPT hutch and transfer pipe under vacuum and leak checked.
TÜV test Nov 21st 9-12 am, rehearsal today 10-12 hrs.
Back

## Slide 5: Optical lasers
SASE 1:
100kHz set point commissioning ongoing.
suffered damage in power amplifier during tests -> delay t. b. d.
SASE 3:
installation and commissioning ongoing, slow but steady progress.
SASE 2:
ca. 3 months behind compared with SASE 3
cabling finished
Laser interlock system installation / testing in progress
still open infrastructure issues, main one being a problem with stability of inner door, which could prevent the interlock from working. TS contacted.
Back

## Slide 6: Detectors – DET
1st AGIPD SPB
  Successful testing of the control system after Karabo upgrade
  Loop 5 updates including testing finished (interlock for big gate valve were added to the loop)
  Ramping up/down temp. using chiller Beckhoff device was successfully tested  it will simplify ooling procedure
LPD at FXE
  Two switch boxes were received from RAL.
  The faulty switch boxes were replaced with the new ones.
    Testing the power cards of the super-modules with an external 48V bench supply to ensure the output current in the expected range.
    Checking and repairing some of the power cables and connectors between the switch box and the power cards.

## Slide 7: Detectors – DET
2nd AGIPD MID
  Installation of the cooling systems for AGIPD is in progress
  AGIPD interlocks are defined we need loop11 to test it
  Vacuum system installed  achieved pressure 1,1 x 10-6 mbar
  Configuration of MPODs for AGIPD detector finished and tested
  All LV power cables were re-tested before they are connected to the detector  work on connection of the cables to the detector ongoing
  Communication with μController and MFPGAs via Karabo control was successfully established for half 2
  Issues:
    Cooling system readiness
    Loop11 (AGIPD interlock & chiller control)  some cables are missing
    Current vacuum system will not ensure a safe operation of the cold detector  temporary solution is found, but it needs to be implemented
    Solution for removal of the front blank flange and installation of the bellow is not yet available  modification of the existing tool is needed
DSSC
  Integration in HERA South
    PLC crate installation finalized. Vacuum PLC has problems with the connection to the scroll pump, will be checked next week

## Slide 8: Detectors – DET
  Motion system PLC testing started today (Thu)
  Missing limit switches alignment tool, drawings were provided too late from DESY  asked AGIPD if we can borrow the tool from them
FastCCD
  Cooling tested, better thermal connection required
  Coupling improved, sensor reseated, to be the tested Thursday/Friday
    Chiller to be retrofitted with stainless steel parts on Friday
  First image readout on Monday
    Activities on beam-line required some mechanical modifications to FastCCD support and delayed further tests.
    Performance tests to be started as soon as cooling is back
    FastCCD online corrections to be enabled and installed by CAS
    Request sent to ITDM to add us to SCS data aggregator
JUNGFRAU
  Beamtime on 19.11. offered by FXE
    Test gain calibration procedure
  Replacement parts for the defective JUNGFRAU 1M at FXE are on the way from PSI
Back

## Slide 9: AE
SQS:
Updates related to interloop communication ongoing
Schedule test for MPS on Tuesday
Loop 12 – running, must be commissioned
SCS:
loop 04 is still running without 3 crates, working on it.
loop 07 – supporting issues, work ongoing
SPB:
NKB coordinated motion – was reported that it’s unstable but the encoder resolution factor should be set properly. AE has explained it and during some spare time this has to be modified in order to get good performance, as in other instruments.
Loop 5 – updated with new interlock definition

## Slide 10: AE
MID
Patch panels missing in MID and HED – they will be needed in the future for better maintenance.
Loop 11 – Issues with missing cables in the Hutch. Information from MID still pending on what should and shouldn’t be connected to the loop. Interlock definition has been already received.
Loop 7 – same issues are appearing.
Continuing with commissioning of the first loops’ components.
PPL
SASE3 PLC – to be deployed today.
SASE2 PLC – provided, working on the firmware.
DCSS
Motion system was set up – tests ongoing.
GENERAL POINT FOR ALL GROUPS
The Beckhoff crates’ system requires of two people to take out and pull in the modules in the racks. That means, people who are not authorized from AE to take out or put in these modules, should not do it. It’s highly probable that connectors are damaged. Always, refer to someone from AE who can give you support or consent on the task.
Back