## Slide 1: Joint Operation & Readiness meeting
Report Operation / Photon Run Coordinator
Report Dispatch/PSPO
Instruments
  FXE
  SPB/SFX
  SCS
  SQS
  MID
  HED
Beam transport
  Vacuum
  X-ray optics
  Photon diagnostics
Optical lasers
Detectors
Electronics
  AE
  EETF
CAS
ITDM
SRP  Separate slides
Technical services
AoB

## Slide 2: SPB/SFX
During last two weeks we worked on the KB mirrors alignment. We made significant progress and got much smaller focus in compare to recent user beamtimes. To be continued.
Upcoming timing tool beamtime (the first week of December): a) Gotthard detectors for timing tool (EHC_PAM) - Module 1: Controller unreachable, not solved so far - Module 2: Controller working and configured, receiver not getting data. ITDM checked network, all looks good. b) Optical setup for timing tool studies is pre-aligned using the alignment laser. The PP-laser required for further work.
Back

## Slide 3: Dispatch meeting report / PSPO
Remind about shutdown activities:
  Comments
  Collisions/Problems you see
  New activities of forgotten items where support groups are needed
  Let PSPO know!
Back

## Slide 4: XPD
XGM
  XTD10-XGM: full range gauge died last Friday (16.11.)  vacuum intervention in tunnel (no beam in north branch for ~3h)  slow signal ok, but pulse-resolved signal has several issues: 4 amplifiers died, spares (new !) have noise, and addl. ringing on signal
  XGMs at XTD1+2+9 and SCS online
  XTD6-XGM set offline until shutdown or HED-request
  Initial XGM-operation training for SCS beamline staff (more trainings to come)
Wavefront
  Test with Hartmann-style Imagine Optics sensor at SCS: proof-of-operation, but >50% less time than planned
PES
  Second Hard X-ray review meeting completed
  SA3 22.11.18: PES data taking concurrent with SA3 mono (calibration vs. undulator settings at various photon energies)
Kmono
  SASE2: vertical undulator offset measurement (w/ S.Karabekyan and L.Fröhlich). ~8 minutes / scan / segment. Automated DOOCS scan and fitting, 10 bunches. Finding maximum is possible even with compression. Procedure also ok for SASE1; at SASE3 to be tested (Kmono uses SA3 harmonics)
  https://ttfinfo.desy.de/XFELelog/show.jsp?dir=/2018/47/21.11_a&pos=2018-11-21T18:08:54
First 6 segments have offsets of 200 to 500 µm. Segments 12+21 are well centered.
Back

## Slide 5: HED
OPT TÜV test passed, cabling ongoing (should be finished within 1 more week)
OPT hutch: installations ongoing until end-Nov to get in shape for first beam
EXP hutch: foundations for rails in IA2 are aligned (MEA2) and epoxy-grouted
Control room: monitors have been ordered, number of computers agreed
Back

## Slide 6: Optical lasers
SASE 1: - high power amplifier alignment (incl. 100kHz) finished. - due to set-back from damage, there is too little time left to commission the 100kHz NOPA before the Timing Tool week. - the PP-laser will now be prepared for 1.1MHz NOPA operation - beam will be passed to instruments at the end of next week SASE 2: - PP-laser layout mask is in preparation, will be placed next week - Laser interlock testing is still under way SASE 3: - high power amplifier alignment (incl. 100kHz) is continuing. - NOPA seed path alignment is continuing - no set back so far
Back

## Slide 7: Detectors – DET
1st AGIPD SPB
  Testing/debugging MDL devices (cooling and power procedure)
  Preparation for AGIPD Commissioning (24-25.11) --> plan agreed with SPB
LPD at FXE
  Nov 15,2018: The LPD was powered up and connected to FEM via Karabo.
    There are 13 modules out of 16 that are working fine.
      Q2M2 : some cooling leakage happened in the past and the module was removed physically.
      Q3M3 : The FEM was failed to connect in the past and the problem is still exist.
      Q4M3 : shows a fault in the super-module after replacing the switch box B.
        After taking a dark image, data variation from tile to tile were noticed and some tiles were saturated.
  Still investigating the problem with the Q4M3.
    The test on the switch box B shows that the output voltages are correct.
    Most probably there is a failure with the power card of the this super-module.

## Slide 8: Detectors – DET
2nd AGIPD MID
  Cooling system for e-boxes operational  first tests are successful
  Low-level components for AGIPD interlocks (part of loop 11) successfully tested (except Julabo chiller)
  Loop 11 was updated with interlock conditions -> to be tested
  All power, control and DAQ cables were connected to the detector
  Electronic boards outside vacuum were successfully powered up
  Control system of the detector is operational
  C&C system: uTCA crate was installed in the experimental hutch and successfully used to trigger AGIPD detector.
  AGIPD DAQ is operational --. first test images were taken with AGIPD @ MID
  Issues:
    Cooling system for in-vacuum part of detector is not yet ready (damaged connector  the pump has to be replaced)
    Current vacuum system will not ensure a safe operation of the cold detector  temporary solution is found, but it needs to be implemented
DSSC
  Motion system under test, progress done.
    Need to fix possible quadrant configuration by next week, since it will not be possible to fully commission the system.

## Slide 9: Detectors – DET
  Limit switches alignment tool borrowed from DESY, system aligned to the best possible.
  Vessel will be closed to confirm that the leak at the lateral flanges is cured after the flange surface has been reworked. Vacuum PLC tests will be finalised.
  Quadrant mounting tool missing some pieces now available.
  Work on external interlock definition (in addition to SIB) started.
FastCCD
  Interlock
    Installed but needs to be commissioned
  Cooling
    Could not stabilize the camera temperature after thermal grease application
    Working hypothesis is bad connection between cold finger and camera caused by shrinking or bad alignment
    Spring loaded mechanism to keep the cold finger in place is being manufactured, installation foreseen on Friday
  DAQ
    Managed to correctly set Bias voltage after reseating clock fiber optics transceiver
    Acquisition hangs up after a few frames due to malformed data
    After a long debugging session with LBNL we concluded that the most probable cause of the issue is the connection with the data fiber optics transceivers
    LBNL will send new transceivers and pigtail connectors next week
    Commissioning to resume this weekend after removing electronics boards from their housings
Back

## Slide 10: AE
SQS:
Loop 04 was updated - pico motors were added, but not tested yet.
Agreed with Patrik G. to work on loop 12 motors next week.
Working with Beckhoff on getting support for servo drives setup.
SCS:
Loop 04 was updated. Remain problems with BKF00320 - it will be checked this week and hopefully installed.
Loop 07 was updated to solve some encoder/motor problems, and include new interlocks (only inside loop).
BKF00206 modified to get scroll pumps running properly.
MID:
Loop 4: Working on issues reproduced in Technosoft, test set up in the lab has been set for this purpose.
Loop 11: initial version deployed, some electrical issues solved. Updated version with interlocks already running in the PLC
DCSS
Tests ongoing. Floating position for motion solved. Homing procedure can be worked around until we can have new framework release.
AE-Framework development
Better test system set in the lab.
Next framework release during these days (1.32.0). It will be agreed with CAS to test it with Karabo.
microTCA
MicroTCA crate in MID experiment hutch install and working – tested with AGIPD detector.
Back