## Slide 1: Insight on Failures of MTCA.4 Systems installed in FLASH and EuXFEL Tunnels
2019 MTCA.4 workshop Hamburg, 05.12.2019
Julien Branlard and Christian Schmidt
for the LLRF team

## Slide 2: AGENDA
INTRODUCTION: FLASH and XFEL
RADIATION and SHIELDING
FAULTS in MTCA.4 LLRF SYSTEMS
OUTLOOK
© alamy

## Slide 3: FLASH and XFEL
3.4 km
FLASH operated with MTCA.4 since 2013
XFEL operated with MTCA.4 since 2017

## Slide 4: MTCA.4 LLRF components* in the tunnels
Only components installed in tunnels are listed
Other components are useful as witness outside of radiation prone environments
 | XFEL | FLASH
ADCs | ~ 350 | 16
DWCs | ~ 350 | 14
CPUs | ~ 55 | 4
MCHs | ~ 55 | 4
TIMER | ~ 55 | 4
DAMC02 | ~ 30 | 2
TCK7 | ~ 55 | 3
PM | ~ 100 | 4
others | ~ 20 | 6
TOTAL | ~ 1070 | 57
* more in other subsystems (technical interlock, special diagnostics, etc…)

## Slide 5: Measuring Radiation Shielding Conditions

## Slide 6: How is radiation measured
Pandora
  FLASH/ XFEL installed in the tunnel
  Measure gamma and neutrons
  Sparse measurement spots
  Mounted on remote controlled robot (XFEL)
RadFet
  Mounted on racks
  Measure integrated radiation
RadCon
  Gamma counters
  Installed as USB stick on MTCA CPUs
  Fast “live” read outs
MARWIN: Mobiler und Autonomer Roboter für Wartung und INspektion

## Slide 7: How is radiation measured
Pandora
  FLASH/ XFEL installed in the tunnel
  Measure gamma and neutrons
  Sparse measurement spots
  Mounted on remote controlled robot (XFEL)
RadFet
  Mounted on racks
  Measure integrated radiation
RadCon
  Gamma counters
  Installed as USB stick on MTCA CPUs
  Fast “live” read outs
Dose rate [Sv/h]
Dose rate [Sv/h]

## Slide 8: How is radiation measured
Pandora
  FLASH/ XFEL installed in the tunnel
  Measure gamma and neutrons
  Sparse measurements
  Mounted on remote controlled robot (XFEL)
RadFet
  Mounted on racks
  Measure integrated radiation
RadCon
  Gamma counters
  Installed as USB stick on MTCA.4 CPUs
  Fast “live” read outs
MTCA crate
external stand-alone sensor
Sensor mounted on FMC
Sensor mounted in dosimon boxes

## Slide 9: How is radiation measured
Pandora
  FLASH/ XFEL installed in the tunnel
  Measure gamma and neutrons
  Sparse measurements
  Mounted on remote controlled robot (XFEL)
RadFet
  Mounted on racks
  Measure integrated radiation
RadCon
  Gamma counters
  Installed as USB stick on MTCA.4 CPUs
  Fast “live” read outs
Courtesy N. Walker

## Slide 10: How is radiation measured
Pandora
  FLASH/ XFEL installed in the tunnel
  Measure gamma and neutrons
  Sparse measurements
  Mounted on remote controlled robot (XFEL)
RadFet
  Mounted on racks
  Measure integrated radiation
RadCon
  Gamma counters
  Installed as USB stick on MTCA.4 CPUs
  Fast “live” read outs

## Slide 11: How is radiation measured
Pandora
  FLASH/ XFEL installed in the tunnel
  Measure gamma and neutrons
  Sparse measurements
  Mounted on remote controlled robot (XFEL)
RadFet
  Mounted on racks
  Measure integrated radiation
RadCon
  Gamma counters
  Installed as USB stick on MTCA.4 CPUs
  Fast “live” read outs
Machine ramp down/up
Single RF station gradient adjustment

## Slide 12: Shielding conditions at FLASH
Primary intention: testing electronics in radiation environment for XFEL operation conditions
  Drawback is accessibility and maintainability
  Benefit of shorter cable length (latency and less prone to environmental conditions)
Gun / ACC1 / ACC39 rack shielding with sliding door in front and back
ACC23 rack shield enclosure without doors

## Slide 13: Shielding conditions at XFEL
Inside the main tunnel : heavy duty concrete blocks
Inside the injector : boron-based shielding identical to FLASH

## Slide 14: Gradient induced radiation (i.e. not beam-induced)
Change in station voltage wrt average
Radiation readout (RadFet estimated daily dose rates)
shutdowns
shutdowns
Lower gradient
(11 GeV run
instead of 14 GeV)
Lower gradient
(11 GeV run
instead of 14 GeV)
Courtesy N. Walker

## Slide 15: Faults in MTCA.4 systems installed in FLASH and XFEL

## Slide 16: SSD failures at XFEL
SSD exchange
  2017 : 15
  2018 : 72
  2019 : 22 (Jan – Apr)
  Not a single failure outside of XTIN, XTL, XTDs since 2017
  Not so many failures in SASE tunnels anymore, but LLRF stations and BC1
Period 2018-02-01 to 2019-04-01
Courtesy Tim. Wilksen

## Slide 17: SSD failures at FLASH
Cryomodule
Beam direction
Gun / ACC1/39 Rack Equipment, GUN crate (2U left side has been moved down to increase distance to beam pipe and CM)
Nb of failures decreased in 2019 , especially for the GUN
  1 BACCA installation in ACC139 Rack
  2 Change of gun crate inside rack
1
2

## Slide 18: SEUs (ADCs only) at XFEL
Single event upsets
shutdown
power outage
shutdown
crate / board reboot
A19
A12
A22
A23
A6

## Slide 19: SEUs (ADCs only) at XFEL
6 boards x 2 crates (master & slave) x 25 RF station = 300 boards
Rate is moderate: 0.15 SEU / RF station / day (12 boards)
Usual suspects: A6, A9, A12, A15-16, A19, A22, A23

## Slide 20: SEUs (ADCs only) at XFEL
Correlation between radiation and failures
  2 months observation period
  Correlation with radiation hits (RadCon) readouts
  LLRF firmware now has single SEU correction block

## Slide 21: PCIe failures
PCIe link to an ADC board disappears during operation
  Not always the same board
  Happens on LLRF crates only (fw? most heavily equipped crates ?)
Diagnostics
  Seems unrelated to radiation conditions
  Coincidentally observed on current consumption
    Current drops (cause or consequence ?)
    PhD devoted to “self healing” systems
  Currently not handled properly by server (work in progress)
Impact of new NAT firmware* ?
  PCIe configuration was changed in new FW
  * New FW  backplane manager support + bug fix in handling slot 12 at reboot
  also, PS redundancy issues were fixed currently under test, installation planned for Summer 2020
failure
board reboot
Board current consumption read from power supply

## Slide 22: Power glitch affecting 1kW power supplies
(*) for the XFEL/FLASH setup
Problem discovered a couple of years ago:
  If power glitch shorter than ~30 msec  crate runs without interruption
  If power glitch longer than ~100 msec  crate stops and restarts
  If power glitch around 50-80 msec  crate sometime hangs and can’t be recovered remotely (power cycle in-situ required)
The higher the load, the shorter the critical time (fans speed, nb. boards etc…)
Affected all MTCA users at DESY
MSK developed test stand and communicated with Wiener
Problem was investigated at Wiener  new firmware MTCA4PS-v1-08-5136-0.hpm (*)
With new firmware, problem could NOT be reproduced
FW update of all 1kW power supplies summer and winter 2019
See Mariusz Grecki’s poster

## Slide 23: OUTLOOK
XFEL (summer 2020)
Operation with backplane manager (MCH-RTM)
Operation with full redundancy (2x 1kW PS)
FLASH
No hardware failure, no MTCA.4 components exchanged because of issues so far
New FW with SEU correction (winter 2019)
Installation of RadCon (winter 2019)
PCIe link failures under investigation (currently dominating issue in terms of LLRF system down time at XFEL)
PhD student to look into HW failure signature and and possible mitigation strategies
More tools to keep detect radiation bursts and draw a radiation map

## Slide 24
Julien Branlard
MSK
julien.branlard@desy.de
+49 (0) 40 8998 1599
Thank you
Christian Schmidt
MSK
Christian.schmidt@desy.de
+49 (0) 40 8998 1869