## Slide 1: BKR training for X-ray Gas Monitors at European XFEL
Theophilos Maltezopoulos
X-ray Photon Diagnostics (XPD)
Hamburg, 19.03.2019
Basic concept of XGM
24/7 operation
Quick fixes

## Slide 2
X-ray Gas Monitors (XGM) setup
Upstream
Downstream
X-ray beam
XGMDh
XGMDv
HAMPh
HAMPv
XGM@XTD10
Single-shot non-invasive pulse energy measurement
and average non-invasive beam position monitoring
Users
Operators
Developed at DESY
in the group of Kai Tiedtke

## Slide 3
XGMDh electrons (4 ns FWHM)
HAMPv ions (100 ns FWHM)
Vacuum
Ions
Electrons
XGMD
HAMP
i1
i2
HVi mesh
HVe mesh
e1
e2
HVe
X-ray beam
Vacuum
X-ray beam
HVrep
Ions
HVHAMP
MP stage
Electrons
h1
h2
Averaged absolutely
calibrated ion signal
↓
Photon flux with
absolute measurement
uncertainty of 7-10 %
Pulse resolved
electron signal
Pulse resolved
ion signal
Software average in timescale of ion current → Cross calibration

## Slide 4: XGM equations
Number of ions created per pulse in the XGMD:
Nph number of photons per pulse
σph total photoionization cross section
zXGMD length of Faraday cup (27.8 cm)
patom target gas pressure
k Boltzmann constant T temperature
Ion current measured by Faraday cup of the XGMD:
qph ion mean charge
TNi transmission of Ni mesh in front of the Faraday cup (80%)
Npulses number of pulses per train
Rrep train repetition rate (10 Hz)
e elementary charge
Courtesy from A. Sorokin and K. Tiedtke

## Slide 5: Cross sections
SASE1 / SASE2:
3 - 24 keV
SASE3:
250 eV - 3 keV
FLASH:
14 eV – 310 eV

## Slide 6: Locations of XGMs in the tunnels
XTD6
XTD9
XTD 2
XGM in XTD6
XGM in XTD10
XGM in XTD2
XGM in XTD1
XGM in XTD9
SASE2
XTD1
SASE1
SASE3
XTD3
XTD5
XTD7
XTD 10
XTD 4
MID
HED
SPB
FXE
SQS
SCS
XTD8
XGM at SCS

## Slide 7: XGM installation
Gas supply (N2, Kr, Ne, and Xe)
Differential pumping up- and downstream (vacuum group)
Vacuum and gas controls under Karabo (photon system control)
XGM measurement, controls, and data acquisition in DOOCS (accelerator system control)
DOOCS to Karabo bridge
Data acquisition in Karabo
XGM@XTD2

## Slide 8: 24/7 operation
0 bunches
2700 bunches
No lasing
Full saturation
SASE search with HAMP signal
100 kHz
4.5 MHz
6 keV
24 keV
Simultaneous operation SASE1, SASE2, and SASE3
Fresh bunch …
Alternating trains pattern …
Number of bunches
Repetition rate
SASE1 photon energy
Pulse energy
XGM has no “official” 24/7 support !!!

## Slide 9: DOOCS

## Slide 10
27.07.2018
SASE1 / XTD2 XGM
SASE3 / XTD10 XGM
Fresh bunch operation with old SASE viewers

## Slide 11
08.03.2019
SASE1 XTD2 XGM with new viewer
Average
Single-shot ADC data
Average ion current
???
???

## Slide 12
HAMP operator
Set needed HV
Max signal level !!!

## Slide 13: Photonflux overview
All green?
Fast signal calibrated to slow?
… only without fresh bunch !!!
Status and errors

## Slide 14: Keithley range
Proper Keithley range selected?

## Slide 15: DOOCS server restart
Restart server,
or “Stop and Offline”
and “Start and Online”

## Slide 16
15.06.2018
0.134 nm, 60 bunches, slow ion signal
both XGMs with Xe, Run 149
Slope: 0.97 (beamline losses)
Averaged pulse energy: XTD2 versus XTD9 XGM

## Slide 17
Slope: 3.8
Slope: 2.5
Distance SASE1 source to XGM@XTD2: 209 m
Distance SASE1 source to XGM@XTD9: 878 m, therefore: 0.4 mm / 878 m correspond to 0.5 mrad
15.06.2018
0.134 nm, 60 bunches, slow ion signal
both XGMs with Xe, Run 149
Averaged beam position: XTD2 versus XTD9 XGM
Pointing!
Pointing!
Can be used for future active feedback

## Slide 18
Position correlation: XGM versus Imager
24.08.2017
0.136 nm, 60 bunches, slow ion signal
XTD9 XGM with Xe, Imager pop-in type I
Slope: 1

## Slide 19
11.11.2018
1.127 nm, 120 bunches, run r0183,
gas attenuator set to 5.5 % transmission
XTD10 XGM with Kr, SCS XGM with Xe
Single-shot pulse energy: XTD10 versus SCS XGM
Slope: 0.02 (overall transmission)
Correlation coefficient : 0.98

## Slide 20
People@DESY … 15 years of experience!
Kai Tiedtke
Ulf Fini Jastrow
Sören Grunewald
Andrey Sorokin
Susanne Bonfigt
Bican “Jimmy” Yilmaz
Machine operators
Transport group, Survey and alignment group
Grounding: Peter Göttlicher, Narcisse Ngada
Vladimir Rybnikov, Olaf Hensler
Tim Wilksen, Raimund Kammering
Lars Fröhlich
Acknowledgments
People@EuXFEL
Jan Grünert
Marc Planas
Florian Dietrich
Joakim Laksman
Torben Falk
Sonay Sayar
Wolfgang Freund
Andreas Koch
Naresh Kujala
Jia Liu
Johannes Risch
AE
CAS
ITDM
Vacuum group
SCS, SPB, HED
E-plan team
TS
… and many more

## Slide 21: Summary
Basic concept of XGM
Position of XGMs in the tunnels
Simultaneous 24/7 operation of XGMs
Quick fixes
Cross-correlation results
Thank you for your attention !!!
Two new publications soon available:
A. A. Sorokin JSR 2019 “XGM for FELs”
Th. Maltezopoulos JSR 2019 “Operation of XGMs at Eu.XFEL”

## Slide 22
29.06.2017 - 2 bunches at 0.15 nm
From synchrotron measurements:
< 10% for absolute pulse energy
< 20 mm for absolute beam position
24.11.2017
550 mJ, single bunch mode, at 0.136 nm
Absolute measurement uncertainty … … with Xe (XTD2): 39 mJ sigma (7%)
… with Kr (XTD9): 176 mJ sigma (32%)
10.11.2017
350 mJ, 2 bunches, at 0.133 nm
Absolute measurement uncertainty …
… with Xe (XTD2): 20 mJ sigma (6%)
… with Kr (XTD9): 90 mJ sigma (26%)
Change of gas: Xe versus Kr
… gives same pulse energy 

## Slide 23
How to get XGM DAQ data
Single-shot pulse energies
Average beam positions
Average pulse energies
Train IDs for sorting data
Average data
Pulse resolved data

## Slide 24
Interlock : not inhibiting = OK !
Gas type
Emergency shutoff valve (status)
Status of gas supply
Gas needle is DOOCS controlled!
“Vatterfly valves”
XGM Karabo scene