## Slide 1: Application of MTCA.4 with RF Backplane to LLRF and BPM Electronics at SPring-8
H. Maesaka, H. Dewa, T. Fujita, T. Fukui, N. Hosoda, M. Ishii, E. Iwai, M. Masaki, T. Ohshima, S. Takano,
RIKEN SPring-8 Center
Japan Synchrotron Radiataion Research Institute (JASRI)
December 2nd, 2020

## Slide 2: Introduction
SPring-8-II
SPring-8
  Beam Energy: 8 GeV
  Lattice: Double-bend acromat (DBA)
  Natural Emittance: 2.4 nm rad
  # of cells: 44
  User service since 1997.
  Electronics are based on NIM and VME.
SPring-8 Upgrade Project (SPring-8-II)
  Beam Energy: 6 GeV
  Lattice: 5-bend achromat (5BA)
  Natural Emittance: ~100 pm rad
  MTCA.4 for high-speed electronics.
    Low-level RF (LLRF), Beam position monitor (BPM), etc.
  We started R&D of MTCA.4 ~5 years ago and we are already upgrading these electronics.
New 3 GeV light source in Japan will also use MTCA.4 for LLRF and beam diagnostic systems.
SPring-8

## Slide 3: Design Concept
LLRF and BPM electronics must detect acceleration RF signals precisely.
  Acceleration RF frequency: 508.58 MHz
Under-sampling scheme was chosen for RF detection.
  The circuit is simpler than down-conversion scheme.
  Drifts of active RF components can be reduced.
    Mixers, LO generators, etc.
  Sampling frequency (fS): 363.3 MHz = 508.58 x 5 / 7
  Intermediate frequency (fIF): 145.3 MHz = 508.58 x 2 / 7
  Digital down-conversion from IF to IQ baseband in the FPGA on the digitizer.
  Required resolution (10 MHz BW):
    Amplitude: 60 dBc = 1 x 10–3
    Phase: 0.1 degree
MTCA.4 can be one of the best choices for the new system.
  A common digitizer AMC can be used.
  The individual RF front-end circuit for each system can be incorporated into a µRTM.
  The RF backplane provides rigid transmission lines and reduces wiring.
Acc. RF
BPM
High-speed
Digitizer AMC
High-speed
Digitizer AMC
High-speed
Digitizer AMC
LLRF Frontend
µRTM
LLRF Frontend
µRTM
LLRF Frontend
µRTM
High-speed
Digitizer AMC
High-speed
Digitizer AMC
High-speed
Digitizer AMC
BPM Frontend
µRTM
BPM Frontend
µRTM
BPM Frontend
µRTM
RF backplane
508.58 MHz
363.3 MSPS
145.3 MHz
[ns]

## Slide 4: Digitizer AMC
0
25
50
75
100
125
150
175
[MHz]
-20
-40
-60
-80
-100
dBFS
0
20
40
60
80
100
120
[ns]
0
10000
20000
-20000
-10000
ADC x 10 ch.
  370 MSPS max., 16-bit
  ADC16DX370
DAC x 2 ch.
  500 MSPS max., 16-bit
  AD9783
FPGA: Kintex7 (XC7K480T)
On-board RAM: DDR3-SDRAM 1 GiB
ADC performance
  Noise level (182 MHz BW): –69.9 dBFS (No input) –51.7 dBc (508.6 MHz input)
    Larger noise than ADC datasheet for 509 MHz input due to clock jitter.
  Amplitude resolution (10 MHz BW): –64.9 dBc ~ 5.6 x 10–4 (508.6 MHz)
  Phase resolution (10 MHz BW): 0.058 deg. (508.6 MHz)
Raw waveform
508.6 MHz input
363.3 MSPS
145.3 MHz IF
FFT
Made by Mitsubishi Elec. Tokki Systems

## Slide 5: RF/Clock eRTM
The primary clock and LO signal are generated by direct digital synthesizer (DDS).
  AD9914
  Up to 500 MHz.
The two clocks and three RF signals (LO, REF, and CAL) are distributed to µRTMs through the RF backplane.
  Clock: 11-way fan out
  RF: 9-way power divider
All the outputs can be turned on and off.
The DDS frequency and output switches are controlled by using UART through LVDS data lines of the RF backplane.
  Operability confirmed with NAT-MCH-RTM-BM-FPGA.
The MMC (IPMI) also works well with NAT-MCH-PHYS80 and NAT-RPM-AC600.
Made by Candox Systems
RF/Clock eRTM
Phase noise data
Additive jitter is less than 100 fs.
Two-slot wide
(extended RTM)
508.58 MHz
363.3 MHz

## Slide 6: LLRF System
LLRF Frontend µRTM
Digitizer AMC
An acceleration RF signal is generated by an IQ vector modulator on a µRTM.
  Baseband signals are given by DACs on the digitizer.
RF signals from high-power components are directly digitized by ADCs after level adjustment by step attenuators on the µRTM.
IF data from ADCs are digitally down-converted in the FPGA.
The klystron forward signal and the vector sum of the cavity pickup signals are fed back to DAC for the stabilization of the amplitude and phase.
The RF switch is turned off when a machine protection interlock signal is detected.
  Waveforms at such an event can be recorded.
We have replaced the electronics for all the four RF stations with the MTCA.4 system.
ç
Level adjust
Kly FB
Cav Pickup x 8
CPU
Kly Fwd
Kly Bwd
Cav Fwd x 8
Cav Bwd x 8
Interlock
IQ MOD
Kly Drv
Ref
Clock eRTM
EtherCAT
Backplane
SW
Dummy load
HV PS
Cavity
Klystron
Master oscillator
Vacuum
Interlock
Tuner
Motor driver
Level adjust
Tuner FB
Level adjust
Cav FB
RTM (RF FE)
AMC Digitizer
MTCA.4
ADC
DAC
ADC
ADC
LLRF µRTM
Front
Rear
Made by Candox Systems

## Slide 7: LLRF Performance
Stability of the vector sum data
  1-week data
  Amplitude: 1.4 x 10–4
  Phase: 0.017 deg.
  (within the feedback loop)
Waveform data in case of a breakdown in a cavity.
  The backward power from the cavity began to increase about 6 µs before SW off.
  This interlock was triggered by over-voltage of the cavity backward signal.
  These waveforms are quite useful for postmortem analysis.
1 week

## Slide 8: BPM Electronics
Two BPM inputs per board.
  (4-ch. BPM) x 2 + (2 ref. inputs) = 10 inputs
508.58 MHz signals are extracted by SAW band-pass filters (~10 MHz BW).
The signal level is adjusted by step att. and amplifiers.
  Attenuation: 0 – 63 dB, Gain > 40 dB
  Input signal level
    100 mA uniform filling: –6 dBm
    0.1 nC single bunch: –53 dBm
Pilot tone generators for gain calibrations.
FPGA Firmware
  Digital down-conversion
  Beam position calculation
  Four kinds of data.
    COD mode, 10 Hz slow data
    COD mode, 10 kHz fast data
    COD mode, Turn-by-turn data (209 kHz for SPring-8)
    Single-pass mode (turn-by-turn)
We are testing the new BPM electronics with 4 BPMs in the present SPring-8 storage ring.
  The present BPM electronics will be gradually replaced with the MTCA.4 system.
Front
Rear
Made by Candox Systems

## Slide 9: BPM Performance
Raw IF waveform (single bunch)
363 MHz clock
IF: 145 MHz
100 ns
31 µm std.
Single-pass
Horizontal
100 µm/div.
0.55 µm std.
COD (10 kHz fast data)
Horizontal
20 µm/div.
BPM Resolution
Single pass: 22 µm (0.13 nC)
COD (fast data): 0.39 µm (30 mA, 2 kHz BW)
Beam orbit was intentionally shaken for these measurements.
20 µm/div.
Filling pattern changed
20 µm/div.
Filling pattern changed
1 month
Long-term stability
The beam orbit was stabilized by the present BPM system.
Beam position data from the MTCA.4 system was stable within 10 µm for more than 1 month.

## Slide 10: Summary
We have developed MTCA.4 LLRF and BPM electronics for the SPring-8 upgrade project.
  RF backplane is used for RF and clock distribution.
High-speed digitizer AMC
  370 MSPS, 16-bit, 10-ch. ADCs and 2-ch. DACs
  Under-sampling scheme to detect 508.58 MHz RF signals.
RF/Clock distribution eRTM
  DDCs for primary clock and LO generation.
  This eRTM can distribute three RF signals (LO, REF, CAL) and two clock signals.
RF Frontend µRTMs
  LLRF: 9 RF inputs with step attenuators and 1 vector modulator output.
  BPM: 2 BPM inputs with SAW filters, step attenuators, and 40 dB amplifiers.
LLRF system
  RF detection, RF generation, ALC and PLL feedbacks, etc. are working well.
  We completed the replacement of the LLRF system for all the four RF stations.
BPM electronics
  BPM calculations for single-pass and COD modes are successfully implemented.
  Sufficient position resolutions and long-term stability were confirmed.