## Slide 1: ARES Operation Meeting
Summary of week 3
Willi Kuropka, on behalf of the ARES shift crew

## Slide 2: Summary
Week 2
 | Monday 11th January | Tuesday 12th December | Wed. 13th December | Thursday 14th December | Friday 15th December
Achievements/Overview |  | Start-up Investigation if laser position on cathode has an influence on secondary beam | Establish beam up to dipole spectrometer Investigate spectrum of the beam | Investigate secondary beam Quad scans FC and DaMon signals | QE measurement Velocity bunching working point
Difficulties |  |  | Magnet power supply problems |  | 
Notes |  |  |  |  | 

## Slide 3: Summary
Week 3
 | Monday 18th January | Tuesday 19th December | Wed. 20th December | Thursday 21st December | Friday 22nd December
Achievements/Overview | Cathode laser investigation (secondary pulse) Alignment of additional larger apertures | Investigate secondary beam | Investigate secondary beam Also on the cathode laser | Investigate secondary beam Gun solenoid moved TWS1 SOL4 moved Changed cathode laser pulse length | Investigate influence on beam of different laser pulse lengths Gun solenoid realigned
Difficulties | No pre- or post-pulse measured with fast photodiode | Faraday cup configuration | Magnet power supply problems Magnet middle layer server was down | IT problems, some tools are not useable Gun solenoid alignment was not satisfactory | One of the bkr consoles is very slow Still IT difficulties
Notes |  |  |  |  | 

## Slide 4: New cathode laser apertures
500um
800um

## Slide 5: QE measurement
Measurement from 2021.01.15
Issues:
Transmission does not scale linearly
Pulse energy only measurable at maximum output of internal attenuator
QE between 4% and 6%
Assumptions:
Linear scaling of transmission along optical path (e.g. before and behind aperture)
Interpolation/Extrapolation of internal attenuator output (no measurement at 15% int. attenuation)

## Slide 6: Comparison to last QE measurement
Scaling at 15%: (71 +- 0.4)%
Scaling at 20%: (76 +- 4.8)%

## Slide 7: Secondary beam
After spectrometer dipole, for different laser attenuator settings

## Slide 8: Secondary beam
On Faraday cup an DaMon
DaMon signal over laser att. position

## Slide 9: Secondary beam
gun phase scan
FC1
DaMon

## Slide 10: Stability measurements

## Slide 11: Stability measurement
20.01.2021

## Slide 12: Plan for this week
Re-check timing of sec. bunch.
MDI shift (to be clarified)
Emittance measurements
Check alignment of magnets
Velocity bunching studies
Charge stability measurements

## Slide 13: Schedule
Week 4
Date | Shiftleader
25.1. | MKK1, MDI, FS-LA
26.1. | Hannes
27.1. | Frank
28.1. | Willi /MDI - shift
29.1. | Thomas

## Slide 14: Schedule
shutdown
Beam operation