## Slide 1
Power Supply Tests for LUXE Gamma Beam Profiler
21.09.2022
INFN Padova
Sergii Vasiukov

## Slide 2
top view
side view
Output Channels
Sapphire Sensors
Monocrystal
University
Wuppertal
Gamma beam profiler as a load
Detector box
The device contains the three different types of sapphire sensors:
W - “Wuppertal”
U - “University wafers”
M - “Monocrystal”
Beam
HV plugs
4-channel GBP prototype
200-channel GBP device
Assembly of three
detectors in one box.

## Slide 3
1M Ohm
1M Ohm
1M Ohm
1M Ohm
10nF
1M Ohm
HV
1M Ohm
1M Ohm
+
Ch 1
Ch 2
Ch 3
Ch 4
Sapphire
Strips
Metallization
Bias voltage
GND
GND
GND
GND
GND
GND
GND
GND
GND
GND
Gamma beam profiler as a load
Scheme and principle
Beam
Sapphire-trips are non-ohmic contacts
Under ionizing radiation, a small current flows in the circuit at a bias voltage. The current gives information about the integral value of the radiation intensity. The signal in each channel gives the cross-section distribution of the beam intensity in the vertical or horizontal direction.
The goal is to reconstruct the beam profile and parameters.

## Slide 4
The bias voltage from 0 to 1.5 kV
The standard Ethernet or USB connections and protocols
High accuracy and stability of bias voltage
High accuracy and reliability of voltage and current monitors
More ???
The minimum requests list

## Slide 5
A1561H
12 Channel 6 kV/20 µA
Common Floating Return Board
LAN/Ethernet interface
S/N 19804
HV power supply module
Model | Channels Number | Max Output V | Max Output I | Vset/Vmon Resolution | Iset/Imon Resolution
A1561H | 12 | 6000 V | 20 uA | 100 / 10 mV | 1 nA / 50 pA
N1471H | 4 | 5500 V | 20 uA | 100 mV | 0.5 nA / 50 pA
Vmon and Imon are important parameters in our experiment
The module under test
N1471H
4 Ch Reversible 5.5 kV/20 µA
NIM HV Power Supply High Accuracy Module
USB interface
S/N 1021
The firmware version is 2.01 (an old release of 2013) and doesn’t include the current offset setting (ZC Detect and ZC Adjust).
Standard module (for comparison)

## Slide 6
Bias voltage affects the current monitor value. The current has never reached zero or negative value.
A1561H + SY5527LC test (no irradiation)
100 pA
1750 pA
2250 pA
4200 pA
50 pA
Imon Resolution
The current monitor value is close to the advertised accuracy of 50 pA at low bias voltage only.
The current offset correction does not completely solve the problem!
No offset correction

## Slide 7
A1561H
Power supply module – Feature No 1
The module under test
The current slowly decreases even after 10 minutes!
The current reaches a minimum positive value of about a few nano amperes.
The current offset correction does not solve the problem at high bias voltage!
N1471H
The firmware doesn’t include the current offset setting (ZC Detect and ZC Adjust).
Standard module (for comparison)
The current drops down to zero in ~10 sec, after reaching the inset HV value.
Only positive current value was observed.

## Slide 8
Primary module crate for A1561H – Feature No 2
A1561H module
S/N 19804
The combination of HV power supply modules and Power Supply System
shows different current values for the same detector!
SY 5527 LC (Padova)
S/N 22458
The current has never reached zero or negative value.
Only positive
N1471H module S/N 1021
Only positive current value or zero were observed.
Positive and zero
SY 5527 (CERN)
S/N 0398
Both positive and negative current values were observed.
Positive and negative

## Slide 9
Questions list
How to measure the exact and real value of voltage and current?
How to make current and voltage measurements independent of other factors and settings?
What setting can guarantee measurement accuracy regardless of a module, crate, software, etc.?