## Slide 1: Status update Rasta
02.06.22
Rapid Sample Test Apparatus
F. Kramer, S. Keckert, O. Kugeler, J. Knobloch

## Slide 2: Existing Infrastructure
QPR
  High quality data (nResolution, , RRR, …)
  VTS (conflicts with other measurements)
  LHe consumption
  Personnel
  ~ 2 weeks turnaround time
Introduction
Design new test cavity which can be less accurate, but faster and less elaborate

## Slide 3: RaSTA – Rapid Sample Test Apparatus
Complementary device to QPR
Identical Samples
Calorimetric measurement like in QPR
Simpler Infrastructure
  Low Power
  No radiation protection
  Small standalone cryostat
  Independent rapid testing
  Less LHe consumption
  High frequency (4.8 GHz)
  Probably higher parasitic losses
  Less accurate
Introduction

## Slide 4: Concept: “Pillbox cavity” in TM020
introduction
QPR sample as an endplate
Design cavity such that H-field zero crossing is at coaxial gap
Prototype from CF components
  Low cost
  Fast delivery
Bottom assembly identical to QPR
Assembly with special RF-gaskets to create pillbox geometry
Yellow Parts: Niobium
  Sample: Bulk or coated
  Cavity: Coated by Universtät Siegen
H field
H field

## Slide 5: Progress since last meeting
Niobium Coating of cavity parts at Universität Siegen (20 µm)
Clean room assembly with QPR Baseline Sample (SS Flange)
1st cold test unsuccessful due to bad electrical contact of RF-gaskets
Re-assembly in clean room with freshly annealed gaskets with maximal achievable torque
2nd cold test successful
Progress

## Slide 6: Mode Spectrum
Mode Spectrum at Room Temperature as expected
TM020 at 4.797 GHz, Q0=6030
(CST: 4.794 GHz, Q=6068)
Results
PCB with AMR sensors
Superconducting
Too low superconducting Q0
TM020 at 4.81145 GHz, Q0=3.46 e5
CST: Q=1.4 e6 (Cu in anomalous limit or RRR 31)
difficult to measure
TM010
TM110
TM020
TM210
TM310

## Slide 7: Superconducting Transition
Quality factor is continuously monitored with VNA during cooldown
Plot Q versus sample temperature
Due to bad thermal contact of sample to LHe bath the cavity becomes sc before sample
Estimation of RRR
Quality factor just before cavity becomes sc: Q=26630
Results

## Slide 8: Dissipated Power
Quadratic increase with field
Higher losses at higher temperatures
Field in cavity is calculated using forward power
Results

## Slide 9: Surface Resistance
RF system is not completed yet
No accurate power measurements possible at this state
Field in cavity is calculated using forward power
Large errors, especially at low fields
Right order of magnitude
Large parasitic losses
Results
T [K] | expected [µ] | measure [µ] | [µ]
2.1 | 0.2 | 23.51.4 | 23.3
3 | 1.9 | 23.81.2 | 21.9
4 | 6.2 | 27.8 | 21.6
4.5 | 9.0 | 30.50.6 | 21.5
T [K] |  |  | 
2.1 | 0.2 |  | 23.3
3 | 1.9 |  | 21.9
4 | 6.2 |  | 21.6
4.5 | 9.0 |  | 21.5

## Slide 10: Conclusion
Successful Nb coating at Universität Siegen
Lesson learned with RF gaskets
Successful Cold test: Proof of principle
  To low Q0
  To low
  Inaccurate power measurements
  Field in cavity
  Calorimetric measurement possible
  Surface resistance in right order of magnitude
  High parasitic losses
Cooperation on simulations with W. Ackermann (TU Da.)
Conclusion & Outlook
Outlook
Adjust input coupling
Build complete RF system
Test QPR sample with coated flange
Design RaSTA 2.0
  Smaller outer dimensions
  Switch from copper gaskets to indium?
  Switch to detachable samples in the future?

## Slide 11: Surface resistance vs B from Pickup
Results

## Slide 12: Versus forward power
Results

## Slide 13: Simulations W. Ackermann
Results
Simulation
Measurement

## Slide 14: CST parameters
5 W input power available with low effort
  no radiation protection needed
55 mW should be sufficient for calorimetric RS measurements
Q0 dominated by copper gaskets
  Acceptable for first test
  Gaskets can be coated later
  Demountable parts for ‘easy’ cleaning, coating, tests …
Rasta parameters
Pamplifier | 5 W | 90 W
U | 0.31 mJ | 5.6 mJ
Psample | 55 mW | 1.0 W
Pcavity | 157 mW | 2.8 W
Vmax | 0.7 kV | 13.3 kV
Q0 |  | 
Bmax | 3.2 mT | 13.7 mT
Emax | 1.6 MV/m | 6.9 MV/m
Pamplifier | 5 W | 90 W
U | 0.31 mJ | 5.6 mJ
Psample | 55 mW | 1.0 W
Pcavity | 157 mW | 2.8 W
Vmax | 0.7 kV | 13.3 kV
Q0 |  | 
Bmax | 3.2 mT | 13.7 mT
Emax | 1.6 MV/m | 6.9 MV/m

## Slide 15
Work packages
WP A | Film production | USI, UHH, HZB, DESY
A.1 | Sample preparation | Nb only?
A.2 | PE-ALD of 1-cell cavity | Up to 3 1-cell cavities, coupon cavity
A.3 | Sputtering studies on flat Nb samples | Flat Nb samples; Nb, NbN, NbTiN, Nb3Sn
A.4 | MgB2 studies | Sputtering on Nb and Cu
A.5 | SC upgrade of RaSTA | 
WP B | Microstructure properties | USI, UHH, BUW, HZDR, HZB
WP C | Field emission and SEY studies | USI, BUW
C.1 | Measurement of parasitic field emission | 
C.2 | SEY measurement | 
WP D | Positron spectroscopy of layered structures | USI, HZDR, UHH
D.1 | Development of an in-situ sputtering chamber | @ pELBE as part of AIDA-II
D.2 | Depth resolved defect analysis at ELBE | 
WP E | Numerical analysis of layered structures | TUDA, USI, BUW, HZB
E.1 | Design optimization, forensic analysis & SC upgrade of RaSTA | 
E.2 | Multipacting analysis based on measured SEY data | 
E.3 | Modelling of an SIS multilayer using a surface impedance representation | 
E.4 | Numerical analysis of the dielectric losses in the insulator of a multilayer device | 
WP F | BCS characterization of samples | UHH, HZB, DESY
F.1 | DC magnetometry of samples | PPMS and AMR magnetometry
F.2 | RF testing of samples | QPR & RaSTA
F.3 | Cavity testing | 1.3 GHz Q0 vs. Eacc with AMR B-mapping

## Slide 16
Milestones and deliverables

## Slide 17
Milestones and deliverables

## Slide 18: QPR sample interface definition
‚naked‘ (dismountable) sample or assembly with DN100CF flange (double sided)
Final height: 85.5 mm
Attention: CERN QPR differs by 3.5 mm in height !
Qpr sample
RF sample surface