## Slide 1
Gravitational wave vertical cryostat
Udai Raj Singh, Jörn Schaffran, and Axel Lindner

## Slide 2
http://arxiv.org/abs/gr-qc/0502054v1

## Slide 3
Discovery of superfluid helium in 1937 by Pyotr Kapitsa, John Allen, and Don Misener in 1937
SF He: fountain effect
Lambda point transition: 2.17 K
boiling suddenly becoming violent for a moment
Superfluid phase below 2.17 K
No gas bubbles in the body of the liquid
SF helium creeping up the inside wall and dripping down as droplets
https://www.scientificamerican.com/article/superfluid-can-climb-walls/

## Slide 4
Applications of SF helium technology
AIP Conf. Proc. 1573, 1214-1221 (2014)
European XFEL accelerator (two phase SF He)

## Slide 5
Applications of SF helium technology
LHC (one phase SF He)

## Slide 6
Phys. Rev. D 100, 092007 (2019)
Applications of SF helium technology
Phys. Rev. D 103, 055017 (2021)
arXiv:2306.09726v1 (2023)
arXiv:2309.07995v1 (2023)
New J. Phys. 19 (2017) 073023
SF He behaving as a qunatum fluid!

## Slide 7
Pressure–temperature phase diagram of helium 4
arXiv:1307.8286 (2013)
Several phases are possible in the heat transfer to helium: superfluid He II, Solid, liquid, and vapor

## Slide 8
SF He
Specific heat of liquid helium at saturated vapor pressure:
Specific heat of He 105 times higher than Cu conductor by unit mass
Lambda-discontinuity in specific heat capacity of liquid helium
In between 1.1 K and 2.17 K, (where as for a metal - )
Change in dominated by rotons and not by phonos
Cu
LHe
Thermal conductivity of liquid helium as functions of and :
Very sensitivity to pressure below 10 K and at higher pressures exhibiting increased thermal conductivity (keeping the bath at 1 atm could be a better solution and not under 10 -30 mbar).
Thermal conductivity over all 103 times that of OFHC copper
Evaporation only from surface, no boiling
Cu
Properties of superfluid helium 4
Steven W. Van Sciver, Helium Cryogenics (2012)

## Slide 9
Properties of superfluid helium 4
Steven W. Van Sciver, Helium Cryogenics (2012)
Specific heat dominated by rotons between 1.1 and 2.1 K and not by phonons
Due to strong temperature dependence, the superfluid component composes about 99% of He II at 1 K.
Phys. Rev. B 103, 104516 (2021)

## Slide 10
Steven W. Van Sciver, Helium Cryogenics (2012)
Russell J. Donnelly, Physics Today 62, 10, 34 (2009)

## Slide 11
General superfluid (SF) helium properties
First sound wave: SF helium has an exceptionally low first sound wave speed of 238 m/s.
Excitations: It also shows unique kinds of excitations, such as quantized vortices, as well as second, third and fourth sound waves.
Density: The low density (ρHe = 145 kg/m3) and effective mass of SF helium contribute to a high zero-point fluctuation (ZPF) amplitude, resulting in significant optomechanical/electromechanical coupling.
Dielectric constant: The dielectric constant of SF Helium is 1.057, indicating its nature as a low-loss dielectric medium. It means that there is minimal light absorption or scattering, resulting in low optical loss.
Viscosity: The absence of viscosity in SF helium results in a very low sound attenuation (acoustic loss proportional to T4 ), which contributes to high mechanical quality factor.
Thermal conductivity: The high thermal conductivity of SF helium prevents any heating issues from RF/optical powers.
Band gap: SF helium does not exhibit linear absorption visible of infrared light, due to its large bandgap (19.9 eV), and the lack of chemical impurities which freeze to the walls of the container.
Tuning tool: SF helium can be employed as a tuning tool for resonators, allowing adjustments to the effective mass and resonance frequency.
In conclusion, the combination of low intrinsic absorption and the large thermal conductivity of superfluid helium contributes to improved RF/optical power handling.

## Slide 12
Steven W. Van Sciver, Helium Cryogenics (2012)
Production of 2 K superfluid helium via isenthalpic expansion
Considering an isenthalpic expansion of helium from 0.1 MPa, 4.2 K to 1.6 kPa corresponding to the saturated vapor pressure at 1.8 K
100% liq.
100% gas

## Slide 13
Joule-Thomson valve
((liq. He )
(liq. He + gas He)
After the expansion of liquid helium (LHe) from high pressure to low pressure through the JT valve, the resulting helium exists in both liquid and gaseous phases, while conserving enthalpy across the valve:
=
Here, , and represent the mass flow rates of the inlet liquid, outlet liquid, and outlet gas, respectively. Meanwhile, , and are the inlet enthalpy of helium, the outlet enthalpy of gaseous helium, and the outlet enthalpy of liquid helium, respectively.
/ = ( / (
In addition, the fraction of mass flow liquefied or yield is defined,
J. Phys.: Conf. Ser. 969 012084 (2008)
/
Furthermore, the mass flow rate through the JT valve is determined by
= . Consequently,

## Slide 14
Lambda plate in the Claudet bath

## Slide 15
1
Developing a new test facility for future cryogenic experiments
General description of HB Box, H1 Box and ACPS in HERA North Hall

## Slide 16
Design and construction of ACPS (ALPS cryo platform subcooler box) in HERA North Hall
ALPS tunnel
location of ACPS
Supplying supercritical helium as well as very high pressurized cold and warm helium gases to three different future experiments
Boundary conditions for helium supply and return in 4.5KF and 5KR circuits: (Tin≈ 4.5 K, Pin≈ 2.5-3.5 bar) and (Tout≈ 6.0 K, Pout≈ 1.0–1.3 bar)
Optimum mass flow rate of helium @4.5 K and 3.5 bar = 30 g/s
Optimal estimation of cooling power @4.5 K and 3.5 bar : 500 W@4.5 K
Helium gas supply and return in 40KF and 80KR circuits: (Tin≈ 40 - 80 K, Pin≈ 13-17 bar) and (Tout≈ 58-66 K, Pout≈ 12–16.6 bar)
Optimum mass flow rate of helium @ 40 K and 14 bar = 75 g/s
Optimal estimation of cooling power @ 40 K and 14 bar : 4400 W
Helium gas supply and return in 300KF and 300 KR circuits: (Tin ≈ 300 K, Pin ≈ 12 - 16 bar) and (Tout ≈ 300 K, Pout ≈ 1.0 – 1.3 bar)

## Slide 17
A detailed view of ACPS in HERA North Hall and transfer lines
H1 Box
ACPS
transfer lines
ALPS quench protection

## Slide 18
Refurbishment of HERA BOX, H1 Box and old transfer lines
HERA Box (existing)
H1 Box (existing)
ACPS subcooler Box (under constrcution)
Currently also working on the refurbishment of old cryogenic boxes and transfer lines in HERA North Hall

## Slide 19
Flow scheme of GW vertical cryostat (GWVC)
ACPS
GWVC

## Slide 20
Flow scheme of GWVC
helium transfer lines (4.5KF, 5KR, 40KF and 80KR)
vacuum vessel
40K/80K thermal shield
sub-atmospheric low pressure manifold 30 mbar and 2 K
U-shape copper liquid-liquid heat exchanger tube
JT valve 1 for producing LHe @ 4.2 K and 1 bar
JT valve 2 for producing SF helium @ 1.8 K and 16 mbar
2K low pressure heat exchanger
high pressure manifold line (300K, 12 – 16bar)
G-10 insulating tube/lambda tube
liquid He bath @ 4.2 K and 1 bar
SF He bath @ 1.8 K and 16 mbar
low pressure manifold (cooldown/warm-up)
cavity support plate
vibration isolation
SRF cavity
PZT actuator
Vacuum pump line
radiation shields
JT valve 3 for producing SF He @ 1.8 K and 16 bar

## Slide 21
U-shape copper HeII - HeII heat exchanger tube
JT valve 2 for producing SF He @ 1.8 K and 16 mbar
2K low pressure heat exchanger
Heat exchangers and JT valves in GWVC
JT valve 1 for producing LHe @ 4.2 K and 1 bar
JT valve 3 for producing SF He @ 1.8 K and 16 bar

## Slide 22
two phase He@ 1 bar & 4.2 K
one phase SF He@1 bar & 1.8 K(no bubbles and high)
two phase SF He@ 16 mbar & 1.8 K(bubbles and less)
Two phase SF (flowing) and one phase helium (static) in GWVC
With one phase helium, it‘s easy to avoid bubble formation and other instabilities like TCOs, TFs, etc.

## Slide 23
Operation modes of GWVC
Warm-up and cool-down of experiments can be conducted in the range of 1.8 K to 300 K, with specific temperature gradients.
The controlled production of superfluid helium using JT valves and heat exchangers will be either single-phase (at 1 bar) or two-phase (at 16 mbar) at 1.8 K.
The steady-state operating temperature will be maintained at 1.8 K.
To cool down the cryostat's thermal shield, helium supply at 40 K and a pressure of 14 bar will be utilized.
Pressure regulation inside the cryostat can be managed by opening and closing control valves.
In the case of a vacuum breakdown, helium will be collected through safety valves.
The cavity can be in direct contact with either single-phase or two-phase superfluid helium.

## Slide 24
Suspensions design
arXiv:2305.01419
PHYSICAL REVIEW ACCELERATORS AND BEAMS 22, 052002 (2019)

## Slide 25
Vapour pressure of helium 4
Minimum achievable temperature here 1.6K/1.7K

## Slide 26
Minimum base temperature of GWVC
10 mbar
100 mbar
Achieving a minimum temperature of 1.6 K to 1.7 K is possible, provided specific conditions are met.
If we could operate the pump @10g/s and minimum inlet pressure 10 mbar/10 kPa.
However, under ideal conditions, the cryostat would operate at 1.8 K and under 16 mbar.

## Slide 27
Warm sub-atmospheric compressor and sub-atmospheric pressure in GHe return line
The mass flow rate is defined as:
where is the density of helium. and are the velocity and cross sectional area, respectively.
The volumertic flow rate is defined as:
/

## Slide 28
Design and thermal analysis of LPCFHEX
Called as low pressure counter flow liquid HeI- vapor HeII heat exchanger (LPCFHEX)
The heat exchanger should ideally exhibit minimal pressure drop, ideally maximum 1 mbar.
Maximum mass flow rate through it ~ 15 -20 g/s
Its primary function is to precool incoming helium, enhancing the liquefaction rate in the cryostat.
Conceptual design and view of external coil of LPCFHEX
AIP Conf. Proc. 1573, 1214–1221 (2014)

## Slide 29
Liquification rate via JT valve
The lower the inlet temperature, the higher the liquefaction rate!!!
A careful design of the LPHEX is required (effectiveness?, Ghe pressure drop?, tube size?.

## Slide 30
Design and thermal analysis of U-shape copper HeII - HeII heat exchanger tube
We still need to design this part.

## Slide 31
Electrical heater power for 2KR@16 mbar GHe return line

## Slide 32
Pipe diameter of 2KR@16 mbar GHe return line
Mean free path of helium molecules in the GHe return line:
: molecular diameter (0.218 nm for helium-4)
= Pressure in pascal
must be sufficiently smaller than the inner diameter of the pump (usually, 5 <d) .
DN80 or DN100 will be fine.

## Slide 33
Heat load budgets for the GWVC
LHe @ 4.2 K & 1.0 bara
SF He @1.8 K and 1 bar
thermal radiation
vacuum vessel@300 K
thermal shield@40 K
MAGO cavity
support structures
penetration, convection, electric connections, etc.
RF power dissipation

## Slide 34
Static RF heat load to SRF cavities in the GWVC
Surface resistance: The Bardeen, Cooper and Schrieffer (BCS) theory of superconductivity can be used to develop expressions for the surface impedance of superconductors on the base of several fundamental material properties like mean free path, coherence length, London penetration depth, and other critical parameters.
Power dissipation in the SRF cavity: At a finite temperature, the unpaired electrons in a superconductor experience the effect of the RF field, resulting in power dissipation on the metallic cavity walls based on their surface resistance and the surface magnetic field intensity.
where is expressed in Kelvin andis expressed in gigahertz [H. Padamsee, RF Superconductivity (2009)].
SCRF Cavity fabrication technology:
The first term represents the BCS contribution, the second term accounts for increased losses resulting from trapped DC magnetic fields, such as Earth's magnetic field. The third term encompasses various sources, including chemical residues on surfaces, foreign material inclusions, condensed gases, and the presence of hydrides and oxide layers.

## Slide 35
The radiative heat transfer from one surface to another is expressed as following (doi:10.18429/JACoW-IPAC2015-WEPTY031):
where is radiative power on surface 2 from surface 1. , , and are temperature, emissivity and area respectively for surface 1 and 2. is view factor that determines proportion of radiation from surface 1 reaching to surface 2. is the Boltzmann constant.
Radiation heat load to GWVC
The empirical formula for calculating the heat flux through the MLI layers sufficiently verified in practice:
Heat flux (W/m2)
= 8.9510-8, = 510-10
: radiation shield layer density (layers/cm)
: number of MLI layers,
= hot side temperature (K)
= cold side temperature (K)
Tm: mean insulation temperature, we take (TH+TC)/2
room temperature (300 K) emissivity of radiation shield (aluminum film emissivity = 0.03)
Nuclear Instruments and Methods in Physics Research A 763 (2014) 701–710 703

## Slide 36
Radiation heat load to GWVC
We need to calculate the area of the 40K, 4K, and 2K cold surface regions that are covered by MLI.
Radiation heat load of 1.55 W/m2 from 300 K to the shield temperature of 40 K with 30 MLI
0.1 W/m2 from the thermal shield at 40 K to 4K and 2K cold masses with 10 MLI.
Tshield = 50 K
Touter vessel = 300 K

## Slide 37
Conduction heat load duo to support structures, RF components, cabling, etc.
For estimating in the case of heat flow within the cold mass, one can use the thermal conductivity (𝑘) and the following formula, where and are the temperatures at the hot end and the cold end respectively, is the length of the bar, and is the cross sectional area of the bar. The unit of is W/K-meter.
Thermal conductivity (and its integral) varies greatly with temperature and with material. A proper choice of materials and thermal intercept strategies for the conduction paths to the cold environment is therefore necessary in the design of the cryostat.

## Slide 38
Radiative heat transfer from 300 K to 40 K = 1.5*3.14*0.5^2*3 = 3.53 W
Radiative heat transfer from 40 K to 4 K = 0.1*3.14*0.5^2*3 = 0.23 W
Different heat loads to GWVC
Heat load due to penetration, convection, electric connections, etc. ?
Heat load due to support structures?
Heat load due to RF power dissipation?
Maximum heat load on GWVC I see here is ~15 W, but the final value from Marc we assume ~30 W.

## Slide 39
ACPS TL3
= 10 g/s, 4.5 K & 3.5 bar
Ideal situation: Liquefaction rate via JT valve and helium return gas
SFHe @1.8 K and 16 mbar to warm compressor
LHe @ 4.2 K & 1.0 bara
SF He @1.8 K and 1 bar
MAGO cavity (heat load = 30 W)
LP CFHEX
JT valve: JT2KF01
JT valve: JT4KF01
= 9.62 g/s
= 0.38 g/s
Enthalpy values of helium supplied through JT4KF01:
Enthalpy of inlet supercritical helium, hi @ (4.2 K, 3.5 bar) = 10.69 J/g
Enthalpy of outlet saturated LHe, @ (4.2 K, 1.0 bar) = 9.90 J/g
Enthalpy of outlet saturated GHe, @ (4.2 K, 1.0 bar)J/g
Enthalpy values of helium supplied through JT2KF01:
Enthalpy of inlet saturated LHe, hi @ (4.2 K, 1 bar) = 9.90 J/g
Enthalpy of outlet saturated SF He, @ (1.8 K, 16 mbar) = 0.84 J/g
Enthalpy of outlet saturated SF He, @ (1.8 K, 16 mbar)J/g
Latent of saturated SF He, @ (1.8 K, 16 mbar) = 22.98 J/g
of SF He = 5.89 g/s
of SF He = 3.72 g/s
Enthalpy values of helium in the bath:
Enthalpy of single phase SF He, @ (1.8 K, 1 bar) = 1.52 J/g
HeII – HeII heat exchanger
Heat absorption = 5.89*22.98 = 135.5 W

## Slide 40
THE GERSEMI VERTICAL CRYOSTAT AT FREIA
arXiv:2103.05265v1

## Slide 41
Cryostat in two parts
Elevating this part to a greater height can help maintain hydraulic pressure during cryostat filling.
Additionally, it can assist in reducing vibrations and minimizing noise.
But it will strongly depend on the fund of the project.

## Slide 42
In which orientation will the cavity be mounted on the cryogenic insert?
Will the MAGO cavity be submerged in SF helium or remain in a vacuum?
What quality factor and resonance frequency would be aimed for?
How do we plan to address the issues of microphonic noise, vibration isolation, flux pinning (very slow cooldown and magnetic shield) and electronic noise?
How can we estimate all potential sources of different heat loads?
Simulation and calculation of various physical parameters for the set-up, such as stress analysis, thermal analysis, etc.
Complete understanding of the flow scheme and translation into a 3D drawing.
Preparation of specifications for the insert, cold compressors, and cryostat.
How many different applications would you like to configure for the cryostat? For instance, applications involving an applied magnetic field or filling the cavity with SF helium.
Specific questions related to the project

## Slide 43
Thank you for your attention!!!

## Slide 44
1908: liquefied Helium and 1913: Nobel prize
Liquefaction of helium and discovery of superconductivity
1911: discovery of superconductivity
H. K. Onnes, Commun. Phys. Lab. 12,120, (1911)
Heike Kamerlingh Onnes (1853 –1926)

## Slide 45
Superconductor-based cables
Superconducting magnet in MRI machine
Maglev trains-400 Km/hr
Quantum computers
Superconducting cavities
Superconducting detectors
Applications of superconductors
Notes: Superconducting magnets for creating magnetic field inside the MRI machine.

## Slide 46
FRESCA2 cryostat at CERN