## Slide 1: R&S calibration kit with Keysight VNA
Akira Miyazaki

## Slide 2: What we (IAS) have in Orsay
KEYSIGHT N5224B
R&S ZV-Z229
IAS professor Bruno had R&S VNA and two calkits in his former laboratory and brought only one calkit to IJCLab

## Slide 3: Incompatible “calkit” data binary files
KEYSIGHT binary
R&S binary
.xkt file
.calkit file
 R&S calibration kit + “file.calkit” is not useful for KEYSIGHT VNA

## Slide 4: Let’s buy KEYSIGHT calkit 

## Slide 5: Let’s buy KEYSIGHT calkit 

## Slide 6: Motivation pour leberité
The authentic calkits exceed our budget
We have less expensive options
  KEYSIGHT firefox calkit: ~ 5 kEUR
  Rosenberger (MPP & UHamburg): ~5 kEUR
  RF Lamda (Aachen): ~5.4 kEUR
  ERAVAN: ~ 7 kEUR
The 3rd party / less expensive option costs anyway a lot
  We could buy receiver chain + RFSoC DAQ board with 6 kEUR!
We already have 3rd party calkit by R&S anyway
 no fundamental change!
We need to reverse engineer the R&S calkit data format
What is “VNA calibration” in the end ????????

## Slide 7: VNA calibration 1 port
To define the reference planes to evaluate S-parameters of DUT not the transmission line around
We connect short/open/load at port 1
3 complex measurement data with Short / Open / Load ()
The actual reflection coefficient of each standards is given ()
Three independent error parameters () can be evaluated at each frequency
https://www.minicircuits.com/app/AN49-016.pdf?srsltid=AfmBOooqx7NqwhZanfjyLFoyBampczlLDHa6g43AwZ7PvChcFP1TlwUJ

## Slide 8: VNA calibration 2 ports
To define the reference planes to evaluate S-parameters of DUT not the transmission line around
We connect Short/Open/Load at port 1 and 2 and then connect Through between 1 & 2
7 complex measurement data for one gender
Critical process before measuring resonant cavities, antennas, with S21 parameter
https://www.minicircuits.com/app/AN49-016.pdf?srsltid=AfmBOooqx7NqwhZanfjyLFoyBampczlLDHa6g43AwZ7PvChcFP1TlwUJ

## Slide 9: Who specifies actual (calibration target)?
 Usually the calkit company!
Two ways
Model based specification
  Circuit model with some parameters over frequency
Real-data based specification
  Factory-calibrated VNA to “calibrate” the individual calibration standards for the customers