## Sheet1
Task | Subtask |  | Duration (# of shifts) |  | Prerequisites | Questions
|  |  | Dedicated | Parasitic
Undulator Beam Transport
| First beam steering |  | 2
| Commision diagnostics |  | 4
| Minimize beam losses |  | 2
| Verify and correct beam optics |  | 2
Electron beam based alignment
| Establish energy steps |  | 1
| Perfrom initial measurement |  | 2
| Verify result and realign intersections |  | 6
| Iterate electron beam based alignment |  | 2
Undulator Effects
| Close undulator one by one and correct field integral feed forward tables |  | 2
23
Photon Beam Transport |  |  |  |  | Photon beam line closed and evacuated Technical commissioning finished Beam Protection System tested (without beam) Personal Interlock finished Radiation Safety approval
Radiation safety measurements with spontaneous radiation (shutter tight?), Personal interlock (?)
| Transmissive Imager | see first spontaneous x-ray photons from a single undulator segment ("first light")
1st SASE search
| Filter box + Transmissive Imager or Filter Box + K-mono Diode | check roughly xray energy (against theoretical value from gap setting)
| 2D-Imager | detection of spontaneous radiation from single segment on 2D imager
Test of absorber system
Test of slit system
GMD
| Diode (at K-mono) | measure angle-integrated pulse energy (photon flux) with diode
| Slits + Diode | find position of maximum intensity (find central ray): fast slit scan (e.g. slits at 1mmx1mm), find maximum signal on photodiode
| Transmissive Imager + Diode pos. | check beam pointing
Beam through double mirror system (beam on FS after mirrors)
| MCP-based detector | first search for SASE (pre-phBBA) |  |  | offset mirrors available ?
2nd SASE search
If beam needs to go to experimental hall: Further radiation tests (interlock?) of shutters towards experimental hall Beam in XTD9
Beam through 600 m of pipe (min. 2 intermediate FS)
Adjustment of focus of mirrors
Test of beam protection system / test of beam absorbers (mass spectrometer)
Test of BPM at beam line end
Beam in experimental hall (Does it have to be focused?)
experiment in XHEXP1
if not:
| K-mono + Diode | single segment (+ref) gap tuning: spectra by changing electron beam energy +/- 1.5% |  |  | wide scan ~ +/-1% range to find emission line, narrower scan ~ +/- 0.5% to acquire spectrum
| K-mono + Diode | repeat for all installed segments |  |  | 1) when using higher order reflexes, reject lower energies with foils 2) Might apply another PBBA method before moving to the next segment (depending on durations)
| K-mono + Diode | calibration scan of K vs. gap for each segment (fixed e-energy) |  |  | measure electron beam energy per shot
| K-mono + Diode | Steepest slope method / gap tuning of 2 adjacent segments: change gap of one segment, take spectra
| K-mono + Imager | optional - Quadrupole kick method: introduce 20µrad kick between these two segments, minimize DK
| K-mono | repeat pairwise segment tuning for all pairs
| K-mono | optional: pairwise tuning of non-neighboring segments
| K-mono | optional: vertical undulator alignment (single segment spectra, dK/K~y^2 )
| K-mono + Diode | Fine adjustment of phase shifters
| MCP-based detector | more search for SASE (post-phBBA): pulse energy and imaging
first lasing at 2 Angstroem
| 2D-Imager | see FEL on 2D imager ("first lasing")
| all XGMDs & XBPMs | commission XGMDs and XBPMs
| XGMD-1 | calibrate FEL pulse energy with XGMD-1 (cross-check with estimate from e-beam energy loss)
| XBPM-1 | calibrate FEL position with XBPM-1
| XGMD-1 + XGMD-2 | check/calibrate attenuators
| XBPM-1 + XBPM-2 | check/calibrate FEL pointing
| Filter Box + Diode (+MCP?) | initial spectral measurements using absorption edge of foils |  |  | K-edges Ge=11103eV(1A), Cr=5998eV (2A), Ni=8333eV, GaAs wafer
| XGMD | Optimize undulator taper (empirically) by intensity
| Pop-in monitor after mirrors | photon assisted offset mirror adjustment (tilt,translation)
transport FEL to XHEXP1
| PES | Commissioning of PES, measure spectrum COG
| Pop-in monitor after distrib. mirrors | photon assisted offset mirror adjustment (tilt,translation)
| all phDIAGS | optimize SASE (vary bunch compression, emittance, RF, BBA, taper,...)
experiment in XHEXP1