## Slide 1: LLRF Capabilities of the European-XFEL
Holger Schlarb
on behalf of LLRF team
Notes: How to edit the title slide
Upper area: Title of your talk, max. 2 rows of the defined size (55 pt)
Lower area (subtitle): Conference/meeting/workshop, location, date, your name and affiliation, max. 4 rows of the defined size (32 pt)
Change the partner logos or add others in the last row.

## Slide 2: Requests: operation flexibility
Questions:
Gradient and phase changes over macro-pulse
Beam loading / beam loading variations
Small charge operation
Related issues:
General stability achievements
Multimode operation
Intra-train Feedbacks systems & Diagnostics

## Slide 3: Parameter overview LLRF system
 | nominal | Min/max
# of RF channels | ~ 3000 | NA
# of LLRF stations 1.3 GHz SRF | 25 | NA
# of LLRF stations 3.9 GHz SRF | 1 | NA
# of LLRF stations 1.3 GHz NRF (RF Gun) | 1 | NA
Field amplitude stability dA/A | 1e-4 | 3e-5/2e-4
Field phase stability | 0.01deg | 0.003deg/ 0.05deg
# of piezo driver/sensors | 800/800 | NA
# MTCA crates | 52 | NA
# RF channels per station (32 cavities) | 118 | NA
Fully automated system with high flexibility and high precision:
  varying beam load: beam current, bunch spacing, bunch pattern
  macro-pulse to macro-pulse structure variations
  changes in gradient, phase, gradient slopes, phase slopes
  incorporate beam based signals (arrival time, compression, energy)
  local and global energy management
  very robust and reliable system with high availability (downtime < 1%)
  ultra-high precision for 20% of RF stations (<<0.01%, <<0.01deg)
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 4: LLRF system layout & functionalities
Reliability / Availability / Maintenance:
  Use industrial standard as bases
  Sufficient diagnostics to identify failures & errors
  Redundancy in DAQ system (Probe/Virtual Probe/Cavity model)
  Redundancy in critical hardware components (Power supply)
  Excellent remote access and controllability
  Handle large number of systems simultaneously
  Data acquisition system for off-line analysis / failure identification
Precision / Flexibility / Automation
  Complex controller with FF, LFF, MIMO-FB, BBF-FB
  System identification, calibration of parameters
  Loop-phase / Frequency Control/ QL control / Cavity phase control
  Quench detection, cavity gradient control and tuning
  Local and global energy management
  Accommodate various beam based FBs
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 5: LLRF system overview
Scheme of LLRF RF controller
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 6: Consequences from FLASH operation
2 semi-distributed uTCA stations (25 RF stations)
Driving Argument Short pickup cables for low drifts (10fs/m/K)
and prevent crosstalk from high power cables
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 7: uTCA LLRF hardware platform
How will a LLRF System look like inside . . . 19“ modules . . .
DCModule (no drifts within the rack)
ISE, DMCS, MSK
LOGModule (Local Oscillators)
ISE, MSK
PIEZOModule (Piezo driver)
Power Supply (redundant)
industry
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 8: Signal integrity: Single channel performance
Short-term stability in a uTCA crate (laboratory) :
Single cavity resolution improved by a factor of 5 to dA/A=2.8E-5.
Signal integrity in uTCA crate achieved Eval board performance.
(using a Power-Entry-Module, Ericsson BMR911483)
Power-Entry-Modules: < 110dB spurious free
Poor Power Supplies : < 80dB SFDR
VS-Scaling : < 120dB SFDR
Testconfguration:
- DWC8300 R1.0, SIS8300_V2, MCH, CPU, uTCA crate
- DRO 1.3GHz (PSI), 1354 MHz LO, 54MHz IF, 81MHz SR
- LO-Generation Module (19“, FL, PM version)
- Offline Matlab non-iq analysis (N=65K), 1MHz Bandwidth
Raw-Data IF Spectrum
(no RF-signal)
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 9: Beam operation using the uTCA-platform
FLASH operation :
Energy stability dE/E=5E-5.
On-crest energy stability :
SR-camera resolution limit
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 10: LLRF Stability
Mixer preserves phase noise
Digital down conversion
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 11: LLRF stability
Reference tracking for mixer drift removal
Mixer phase drifts ~ 0.2/K
Mixer amplitude drifts ~ 0.2%/K
+ dependence on humidity on PCBs
Remark: mixer drift not equal (one PCB, temp.)
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 12: Variation of gradient and phase
Collimator limits energy slope to ~ 3% pkpk

## Slide 13: Multiple – gradient/phase operation
Preparation of software for FLASH 1 & 2 & 3 operation!
FLASH 1
FLASH 2
FLASH 3
Used gradient steps -2% and + 3% with 40us transient time
Setting up time < 10 sec

## Slide 14: Multiple – gradient/phase operation
Preparation of software for FLASH 1 & 2 & 3 operation!
FLASH 1
FLASH 2
FLASH 3
Used gradient steps -2% and + 3% with 40us transient time
Setting up time < 10 sec

## Slide 15: Multiple – gradient/phase operation
Preparation of software for FLASH 1 & 2 & 3 operation!
FLASH 1
FLASH 2
FLASH 3
Used gradient steps -2% and + 3% with 40us transient time
Setting up time < 10 sec
Klystron forward power for
dA/dt ~ 1.5%/20us
Approx. 70% for cavity filling
1%/10us is approx. limit
transient can be improved
(only LLF used) by addition
calc. FF tables
Slow slopes in phase & ampl.
is none issue
Main point:
Appropriate operator interface
Selective Exception handling

## Slide 16: Beam loading and varying beam loading
Gradient flatness
  -> orbit variation across macro-pulse
  -> achievable gradient
Stability
  -> Single bunch loading 1nC: dA/A ~ 1.4e-4
  -> variable pattern requires single bunch correction
Coupler kicks due to varying Pfor/Pref at coupler
Beam loading & high gradient (at quench limit) &
High stability is most challenging operation mode
Fundamental for LLRF system is 9 MHz readout

## Slide 17: Upgrade of LLRF system (2010-2011)
Upgrade of all RF stations using SimconDSP controller
RF control for 3.9GHz
New cabling in injector racks
Upgrade & unified FPGA controller firmware
  Multiple feed forward table (main/beam loading/correction)
  Multiple set-point table (main/beam based correction)
  Model based Multiple In Multiple Out (MIMO) controller
  Charge correction & intra-train beam based feedback
  Exception & Error handling, limiters
  Error and status displays
Feed forward table architecture
Scheme of LLRF RF controller
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 18: Performance of LLRF system after upgrade
ILC studies: 3 MHz, 1.5 nC = 4.5mA XFEL Parameters
All field control applications are turned on:
LFF, MIMO FB, BLC
dA / A < 10-4
dP = 0.011 deg
Acceleration Phase ACC1
BLC 1mA
x charge x rep. rate => 4.5mA
Truncated
Remark: single bunch transient > 1e-4
Forward
1st bunch start
LFF stop
Acceleration Voltage ACC1

## Slide 19: Beam loading and varying beam loading
Beam loading table to short (almost no effect on energy but orbit changes)
9 MHz measurement and 81 MHz DAC output for smoothing of kly input
Signal temporal alignment in LLRF ~ few 10 ns

## Slide 20: Beam loading and varying beam loading

## Slide 21: Performance of LLRF system after upgrade
ILC studies: energy stability / gradient flatness / gradient limit
Minimizing slopes by QL tuning
Important studies for FLASH & XFEL
Impacts orbit variation and orbit slopes
Achievable energy gain
flatness
Energy stability exit of accelerator
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 22: Beam loading and varying beam loading

## Slide 23: Beam loading and varying beam loading
open points:
Set stable SASE operation for 4.5 MHz repetition rate (in conjunction with ILC studies)
Develop interface to new timing system providing the corresponding bunch pattern
Develop tools to deal with rapid varying beam load
Developed at FLASH for operation FLASH1/2/3
Preferable if FLASH is operated within the next 3 years with XFEL bunch pattern

## Slide 24: Layout longitudinal FBs for European XFEL
Gun
I1
3rd
~
~
Laser
BC0
L1
L2
L2
L2
L3
BC1
BC2
L3
L3
COLL
BAM
BPMs
BCM
EO
LLRF
LLRF
LLRF
A/
BAM
BCM
EO
BAM
LLRF
A/
A/
BAM
BCM
SPEC
BAM
LLRF
A/
LLRF
A/
LLRF
A/
OXC
LLRF
A
LLRF
A
BPMs
BPMs
BPMs
uTCA
~
~
BAM
BAM
BAM
Toroid
OXC
BBF Charge
BBF Osc.
BBF BBO
3 feedback loops to stabilize photoinjector laser & charge
BBF Gun
BBF I1
BBF L1
BBF L2
4 feedback loops to stabilize arrival times, compression, shape & energy
BBF L3
1 feedback loop to stabilize final beam energy
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 25: Beam based feedback
Gun
ACC1
3rd
ACC2
ACC3
ACC4
ACC7
~
~
BAM
BCM
BCM
BAM
BAM
Laser
LLRF
LLRF
LLRF
A 
A 
Beam Based Feedbacks:
BAM and BCM after BC2  amplitude and phase in ACC1 and ACC39
BAM and BCM after BC3  amplitude and phase in ACC23
BC2
BC3
R56=180mm
R56=43mm
R56=0.8mm
FLASH 1
FLASH 2
FLASH 3
LLF + BLC + MIMO + BBF at ACC1/ACC23
(act still on setpoint)
< 20fs rms arrival jitter
+ beam based SP correction
< 75fs pkpk
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 26: Beam based feedback
Gun
ACC1
3rd
ACC2
ACC3
ACC4
ACC7
~
~
BAM
BCM
BCM
BAM
BAM
Laser
LLRF
LLRF
LLRF
A 
A 
Beam Based Feedbacks:
BAM and BCM after BC2  amplitude and phase in ACC1 and ACC39
BAM and BCM after BC3  amplitude and phase in ACC23
BC2
BC3
Timing jitter
behind BC2
Voltage
FLASH: 7.0ps/%
2 ps/deg L-band
0.05 ps/ps C=20
Phase
Incoming
compression
factor
C ~5 … 20
Jitter from BC2 largely remains + additive jitter!
Timing jitter
behind BC3
FLASH: 0.9ps/% for dV2/V2 < 2.5e-5 with uTCA => 2fs rms
R56=180mm
R56=43mm
R56=0.8mm
FLASH 1
FLASH 2
FLASH 3
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 27: Beam based feedback: new topology
Gun
ACC1
3rd
ACC2
ACC3
ACC4
ACC7
~
~
BAM
BCM
BCM
BAM
BAM
Laser
LLRF
LLRF
LLRF
A 
A 
BC2
BC3
R56=180mm
R56=43mm
R56=0.8mm
FLASH 1
FLASH 2
FLASH 3
Beam Based Feedbacks:
BAM and BCM after BC2  amplitude and phase in ACC1 and ACC39
but mainly slow variations, systematic & repetitive errors <~20kHz
BAM and BCM after BC3  amplitude and phase in ACC23
BAM after BC2  feed forward drive for normal conduction cavity
Remark to fast long FB:
Only location possible prior to BC2: large R56 at small energy (150MeV)
Typically ~ 50-100 kV @ 1000 x large cavity bandwidth
Short cables, fast processing time, only semi-conducting amplifier
Latency < 1 us
Shoot for FB bandwidth > 100 kHz
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 28: Beam based feedback: new topology
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 29: Beam based feedback: new topology
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 30: Other issues
Detection of soft quench
specs Q0 > 1e10
likely Q0 ~ 2e10
load QL ~ 3e6
soft quenched cavity ~ 2e7 => eats up same power at entire linac!!!
but QL drops 3.0e6 -> 2.6e6 (~ 10%)
Operation recovery
At level dA/A < 1e-4 other passband modes have to be considered
8/9pi mode (instabilities)
7/9pi mode (contributes to dE of beam)
Diagnostics
Arrival time monitors e.g. dynamic range ~ 3 ps

## Slide 31
Thanks for attention
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View