## Slide 1: Can we move the SCS instrument to 2.4 m distance?
20.1.2012
Harald Sinn WP73

## Slide 2: Current Plan (CDR update)
experiment hall
G1, G2, (M4)
M1
M2
M3
M5
M6
horizontal focus
vertical exit slits
Top view
Side view
SCS
281
283.9
300
301
339
341
374
400
417
1.4 m
1.4 m

## Slide 3: Version A: increase angle
experiment hall
G1, G2, (M4)
M1
M2
M3
M5
M6
horizontal focus
vertical exit slits
Top view
Side view
SCS
281
283.9
300
301
339
341
374
400
417
1.4 m
2.4 m

## Slide 4: Change of reflectivity
Energy -10%
due to lower density

## Slide 5: Version B: move mirror
experiment hall
G1, G2, (M4)
M1
M2
M3
M5
M6
horizontal focus
vertical exit slits
Top view
Side view
SCS
281
283.9
300
301
305
341
374
400
417
1.4 m
2.0 m
horizontal focus

## Slide 6: Impact of horizontal focus change
Beam already focused significantly at grating
  40% less damage tolerance at grating
  (91 mJ 54 mJ, B4C coating, 290 eV)
Beam size larger in experimental hall
(for flat M5 e.g. 800 eV: 3 mm 15 mm, vertical size 1.7 mm
Or M5 focusing: hor. 43 um  67 um)

## Slide 7: Conclusions
Shift from 1.4m to 2.0 m possible by shift of distribution mirror. The damage tolerance on the grating reduces by 40% and the horizontal focusing deteriorates by 50%.
Increasing the take-off angle lowers the transmitted upper energy to 1.8 keV (2.4 m) or 2.2 keV (2.0 m)
Whispering gallery also possible (add 0.5 M€)
Safety and transport implications would have to be considered.