Pixel chip testing at DESY
Daniel Pitzl, Alexey Petrukhin DESY
Meeting at PSI, 24/01/2012
- Installation at DESY
- Some test results
- Known issues, open questions

---
Universal pixel test board
psi46
chip
other
adapters
available
Design and firmware
by Beat Meier, PSI
600 V
bias
scope
ADC
trigger
FPGA
memory
USB1
to
laptop
load
FPGA
6 V
power

---
psi46expert software
- C++ class library.
- Written by Peter Trüb (ETH, 2005-2007) for Scientific Linux 32 bit.
- Now compiled with g++ 4.4.5 under Ubuntu 10.10 and SL6 64 bit.
- USB interface required some changes (long → int).

---
Configuration
- Configuration files for test board and readout chip imported from PSI:
  - board name,
  - run in 40 MHz mode,
  - set 26 DACs and Control Registers on the ROC,
  - define timing sequence: reset - cal - trigger – token,
  - read and load calibration and trim values,
  - run in 'psi46expert' or 'takeData' mode

---
psi46 DACs
 13 VIBias_Bus 30
 14 Vbias_sf 10
 15 VoffsetOp 55
 16 VIbiasOp 115
 17 VOffsetR0 120
 18 VIon 115
 19 VIbias_PH 130
 20 Ibias_DAC 122
 21 VIbias_roc 220
 22 VIColOr 100
 23 Vnpix 0
 24 VSumCol 0
 25 Vcal 200
 26 CalDel 125
 27 RangeTemp 0
  1 Vdig 6
  2 Vana 150
  3 Vsf 160
  4 Vcomp 10
  5 Vleak_comp 0
  6 VrgPr 0
  7 VwllPr 35
  8 VrgSh 0
  9 VwllSh 35
 10 VhldDel 130
 11 Vtrim 7
 12 VthrComp 124
253 CtrlReg 0
254 WBC 20

---
psi46 pixel readout chip
4 trim bits
adjustable by programmable DAC, per ROC
programmable register, per pixel

---
psi46 pixel readout chip
adjustable by programmable DAC

---
Some test results

---
CalDel and Time walk
- One pixel.
- Vcal 200 small DAC.
- Moving the threshold shifts the timing:
  - time walk.
noise limit
earlier
BC
later
BC
Vcal too
small

---
Sample and hold timing
250
- One pixel.
- Position of maximum depends on pulse height:
  - time walk.
DAC 150 is compromise
?
200
150
100
50
Vcal

---
GUI
Calibration
 tests
Individual
tests
Parameter
adjusting
D. Pitzl, A. Petrukhin: ROC testing at DESY
Meeting at PSI, 24.01.2012

---
Threshold optimization algorithm
- If comparator thresholds are adjusted with global VthrComp only: spread of thresholds in ROC ~300 e- due to transistor mismatches
- Unify pixel thresholds by 4 trim bits (values from 0 to 15) and scale with Vtrim DAC
- Each trim bit value is set such that Vcal -threshold of the pixel differs least from the selected target threshold in the procedure
- threshold spread reduced to 50(80) e- after the optimization chip without(with) sensor
D. Pitzl, A. Petrukhin: ROC testing at DESY
Meeting at PSI, 24.01.2012

---
Comparator threshold
soft threshold
- One pixel
- Analog pulse height vs threshold and calibrate amplitude.
White region:
  - no signal.
Colored bands are not vertical:
  - time walk.
noise limit
hard threshold
X-ray calib:1 DAC = 65 e
threshold curve

---
Threshold curve
one pixel
vary test pulse amplitude
vary trim bits
TB
1
6
10
15
X-ray calib:
1 DAC = 65 e
threshold broadened by noise
fit by error function
noise: 130 e
preamp
shaper
driver

---
Threshold variation
4160 pixels / cip
the same chip, trimmed:
CMS transistor variations:
threshold spread 290 e
4-bit DAC trimming:
threshold spread 50 e

---
Threshold variation untrimmed
- Vertical scale – threshold in ke
- Chip 0 (no sensor)
- VthrComp 124
10% spread

---
Threshold variation trimmed
- Chip 0 (no sensor)
- VthrComp 112
- Vtrim 104
- 4160 TrimBits set.
- time: 3 min / chip
- 1.7% spread
- 50 e threshold variation!

---
Towards the minimum TrimVcal
TrimVcal=40 TrimVcal=32 TrimVcal=28
Use as low as possible target Vcal: good charge sharing, good
for radiated chip with low charges.
Chip2 (no sensor)

---
Tests with sensor
TrimVcal=80 TrimVcal=70 TrimVcal=60
Underflows: bad fit of SCurves
Chip6 (with sensor)

---
Threshold Scan and Power
- Chip 2 (no sensor)
- Id (mA) vs VthrComp
- Threshold into noise
- Chip consumes 2.5 times more power at low threshold (high VthrComp)
- Chip 2 (no sensor)
- Ia (mA) vs VthrComp
- Small effect at low threshold
- The same trend for Chip 6 (sensor)
Id
Ia
VthrComp
VthrComp

---
Test beam hit map
- 2 GeV e+ test beam
- Chip 6 (sensor)
- 2 dead pixels
- 4 dead regions
- Use 'psi46expert' to test bump bonding

---
Bump Bonding Test Procedure
- Vcal to switch 2 induces a signal in sensor. If bump bond (bb) is present it is seen by preamplifier: missing bb can be identified
- Problem: cross-talk via a parasitic coupling between the calibration voltage line and preamplifier can fake a signal even without bb
- Determine Vcal_Thr2 for the signal injection through the sensor
- Measure Vcal_Thr1 for the parasitic cross-talk (with both switches open)
- |Vcal_Thr1 – Vcal_Thr2| < 5 DAC units → defect bump bonding

---
Bump Bonding Test Results
D Threshold
- Threshold difference: expected defects from beam test are correctly identified by 'psi46expert'
- Peak at D Threshold ~ -20 → good bump bonds. Peak at D Threshold ~ 0 → bad bump bonds

---
Arm Pad, Xtalk
- Three ways to inject charge: Standard (used so far), via 'pad' and through 'Xtalk'
- Different signals can be used for bump bonding test of modules, cross calibration of ROCs and some other purposes ?

---
Arm Results
- Use three different capacitances to inject charge: direct, via sensor, and crosstalk:
  - Q = Cx · Vcal
- One pixel activated.
- Draw PH [ADC] vs calibrate amplitude for large Vcal (450 e/DAC)
- Use sensor type ROCs
- Chips 6, 8 show similar behaviour but PHs are different (different config. parameters for analog gain and offset)
- Blue and Red curves start to saturate later compare to black - ?
Chip 6
Chip 8

---
gain and linear range
- One pixel.
- 2 Vcal ranges (PSI X-ray calibration):
  - CtrlReg 0 or 4,
  - 65±5 e/DAC,
  - 450 e/DAC.
- Linearity for small pulses important for spatial resolution using charge sharing.
- Saturation around 36'000 e (~1.6 MIP).
preamp/shaper
saturation
450 e/DAC
65 e/DAC
linear
threshold at ~3'000 e

---
Linear range vs Vsf
- One pixel
- Analog pulse height vs calibrate amplitude and source follower voltage.
Best linearity in valley.
below
threshold

---
Pixel detector in DESY test beam
2 GeV e+ beam.
- Landau ⊗ Gauss:
  - Fit not perfect:
  - Peak position and width OK for 285 um silicon,
  - noise too large.
- Small peak at small pulse height:
  - Wrong timing due to missing synchronization between clock and beam.

---
        Cluster charge
- 2 GeV e+ beam test
- No magnetic field
- Test pulse gain calibration applied
- Chip 8, all pixels
- Moyal fit to each column
- Expect ~25 ke from 285 μm silicon
- Observe ~8% gain variation across chip8:
  - Test pulse (calibration) problem?
  - Check with X-ray source?

---
Threshold Scan Procedure
- DESY 2 test beam: 2 GeV e+, 5 kHz scintillator trigger, Vbias -90 V
- Chips 6, 8 (sensor), calibrated, trimmed, optimal DAC parameters
- 50 sec runs, 0 – 140 k clusters per run
- Change VthrComp from lowest to optimal DAC units. Large DAC = soft threshold

---
Threshold scan in the test beam
Spec.
DAC ?
Chip 6
Chip 8
 VthrComp
 VthrComp
- Uniforming of thresholds brings more clusters per event
- Different chips show closer behavoiur after the trimming procedure
- Special DAC value found for Chip 8?
cluster/event

---
Known issues

---
Pixel map
Aout [ADC]
- 52 × 80 = 4160 pixel per chip.
- Vcal = 200 DAC
- VthrComp = 80
- Strong pulse height variation:
  - gain?
  - timing?

---
Open questions
- Many optimization procedures use fast reading of FPGA data (scurve data) but it works in TBM emulator mode only
- We can not see UB level in TBM emulator mode. Firmware VersionATB V6.1 (27.08.2010 15:35) ID=22A71BED TS=4C84D636.
  Done so far:
  - 3d scans: clk (LHC clock), sda (DACs programming), rda (token out signal). Takes 2 hours each, ~2 days in total → no UB level
  - TBM DAC 'Dacgain' (UB TBM) and 'Ibias_DAC' (UB ROC to UB TBM) have no effect to UB level
  - No TBM:
  - TBM (gain 4, longer header and trailer):
We need to use bigger time delays between FPGA and USB before data transfer to PC. Fast computers, slow USB-1 ? Example:
  - Data lost at beam test: slow USB, time window was not wide enough for sending of large data blocks. Solved by reducing of data block size sent via USB from 215 to 10.000
Data: 6 : -211 17 126 : -45 15 190 73 -44 189 : 7 8 7 8 7 9
Data: -750 -750 -750 0 374 374 186 : 374 68 500 : -172 64 752 292 -176 764 : 32 -750 -750 0 0 374 -188 186 -188

---
Back up slides

---
Temperature test [Temp]
The CMS pixel readout chip has a built in temperature sensor designed to measure
temperatures in the range [-30°C, +70°C]
Test signal is a LastDac ADC signal (LD) ~ to voltage difference between T sensor (temperature
dependent) and reference (temperature independent)
For better accuracy 8 different windows (reference voltages 400 – 564 mV) of measurement
are defined: one part of T° measurements compared to one V window, another part to another
V window etc. Windows are programmed by register 27 'RangeTemp'
Calibrate output signal to extract the voltage difference from the measured LD ADC value.
The sensor output signal is replaced by 470 mV for calibration. Make a linear fit to the calibration
points differ from the ROC Black level. This fit is used to determine the measured voltage
difference, which is finally added to the known reference V to get a sensor voltage

---
T measurements
PSI46 Specification:

---
Temperature test results
Gain calibration
Measurement
at lab temp
Out of range points
 Voltage difference = Vref – 470 mV
Example chip 2:
Look at ref. Voltage = 425 mV
measure LD = 800 ADC,
=> V difference = 58 mV
Sensor Vtemp = 425 + 58 = 483 mV
T = 0.625 ( Vtemp – Vmid )
Average temperature (°C) measured by different ROCs
Chip 0	Chip 1	Chip 2	Chip 3	Chip 4	Chip 5	Chip 6	Chip 7	Chip 8
17	16	21	19	17	21	20	17	16
Temperature may increase by 3°C during test