## Slide 1: X-Ray Radiation Damage: Sensors
Introduction
Basics on X-ray (Si-SiO2) Damage
Measurement of the relevant Parameters
Impact on Sensors and Sensor Design
Summary
Work supported by:

## Slide 2: Introduction
Bulk vs. surface damage in silicon sensors:
  Si-bulk damage (energetic particles/hadrons):
   Lots of user experience + basic R&D:
  Radiation produces point and cluster defects
    Increase of dark current
    Change of doping (incl. sign inversion) + of shape of electric field in sensor
    Charge losses due to charge trapping by defects
  Many groups work on this topic – vital for particle physics
  Si-SiO2 surface damage (X-rays and low energy electrons):
  Little experience from sensor users + little basic R&D
    Reports (mainly) on increase in leakage current  problems ASICs
  Lots of experience + basic R&D for electronics:
    Change of transistor characteristics
    Leaking capacitors, isolation break-down, and more
Do we have to worry at the XFEL ???

## Slide 3: Introduction
What are the worries ?
  XFEL: a new regime for X-ray science:
    Dose up to 1 GGy non-uniformly over sensor [ = 1 GJ/kg !  ~ 5 MJ in sensor of AGIPD type]
    0, 1 ... 105 photons/( < 100 fs pulse)
    Time between bunches: 220 ns
The worries:
    Breakdown
    Dark current
    Increase of inter-pixel capacitances
    Charge losses
    Slow pulses due to plasma effects (high eh-densities)
    Pile-up
    Change of performance (with time, over sensor area)
I will attempt to address some of these worries

## Slide 4: Basics of Surface (X-ray) Damage
Effects of radiation as function of γ energy:
>1.12 eV: eh-pairs in Si  signal [on average ~ 3.6 eV/(eh-pair)]
> 8.8 eV: eh-pair in SiO2 [on average ~ 18 eV/eh-pair]  positive charges due to low mobility of holes  “surface” radiation effects
>300 keV: crystal damage: point defects (vacancies + interstitials  kinetics)
> 3 MeV: crystal damage: cluster defects
 for 12 keV X-rays only “surface – SiO2” damage
radiation
Properties of thermally grown SiO2 | 
density | 2.27 g/cm3
diel. constant | ~ 3.8
diel. strength | (5 – 10)∙106 V/cm
energy gap | 8.8 eV
spec. heat | 10-3 J/(kg∙K)

## Slide 5: Damage of SiO2 and at Si-SiO2 Interface
Generation of
eh-pairs in SiO2
for 500 nm SiO2: 4∙1016 cm-2 for dose of 1 MGy
Prompt recombi-nation of eh-pairs
Transport of
free carriers
depends on ionization density and E-field
M.R. Shanefield et al.,
IEEE TNS 38(1991)1187

## Slide 6: Damage of SiO2 and at Si-SiO2 Interface
Effects
on sensor
F.B. McLean,
T.R.Oldham
(1987)
Transport of
free carriers
Trapping
at Si-SiO2
interface
Trapping
in SiO2
μelectrons ~ 20 cm2/V∙s  electrons escape
„μholes“ < 10-5 cm2/V∙s  holes trapped in SiO2
or (mainly) at Si-SiO2 interface
[there are O(1015 cm-2) surface states]

## Slide 7: Damage of SiO2 and at Si-SiO2 Interface
Si
~3 nm
SiOX
SiO2
+
+
+
+
-
+
+
+
+
+
+
+
+
-
-
-
+ Oxide trapped charges (Nox):
- mainly positive oxygen-vacancy defects
(one shallow trap  hole transport,
+ one deep trap E’γ @~3 eV) saturation: h-trapping = e- recombination
+ + + + + + + + + + + + + + + + +
+ + + + + + + + + + + + + + + +
+ Border oxide traps (“add” to Nox):
positive E’γ defect can exchange
charge with Si depending on Fermi-
level on time scales 0.01 s - seconds
x x x x x x x x x x x x x x x x x x
X Interface traps (Dit*)):
traps at interface (no barrier !)
dangling bond defects (Pb) – H+ released when h captured:
SiH + H+  (Interface Trap)+ + H2
no. limited by no. of dangling bonds
Mobile ions: not an issue anymore
-
*) Distribution of traps in the Si bandgap
Dit [1/(eV∙cm2)]

## Slide 8: Impact of Oxide + Surface Damage on Sensors
Effects on p+n sensors:
  Accumulation layer
  Lower breakdown voltage
  Increase of inter-pixel capacitance
  Increase of leakage current
  Charge losses close to Si-SiO2
Tasks:
  Measure Nox and Dit vs. dose for different technologies
   Implement results in TCAD sensor simulations
   Optimize sensor design
   Verify by experiment
J.Schwandt

## Slide 9: Irradiation Set-up at F4 (DORIS)
DORIS will have stopped operation by end 2012 !!!
H.Perrey
J.Zhang

## Slide 10: Measurement Techniques + Structures
Irradiations: 10 kGy  1 GGy
(immediately after irradiation not all results reproducible due to short term annealing effects  most of measurements made after 10’@80°C)
- MOS capacitor: C/G-V vs. frequency ….. Nox + Nit
vs. Vmax ……….. border traps
TDRC …………………………… Dit and Nit
- Gate Controlled Diode: C/G-V …………….  Nox + Nit
I/V …………………  Isurface (Dit)
- Strip sensors: I/V  compare to TCAD simulations
C/V  compare to TCAD simulations
TCT  charge losses + surface effects
6 mm x 6 mm test field
- MOS Capacitor (MOS-C)
- Gate Controlled Diode (GCD)
(from different vendors and crystal orientations)

## Slide 11: MOS Capacitor: Principle of Measurement
V
V
V
Ideal
before
irradiation
after
J.Zhang
non depleted bulk
non depleted bulk
non depleted bulk

## Slide 12: MOS Capacitor: VFB-shift vs. X-ray Dose
VFB-shift measures Nox and Nit
  Increase up to 10-100 MGy
  Decrease above 100 MGy ?
  Need additional measurements to separate contribution from Nox and Nit
SiO2
Si depleted
Si non-depleted
Al
H.Perrey
J.Zhang

## Slide 13: MOS Capacitor: VFB-shift and Border Traps
VFB-shift shows strong hysteresis effect
Biasing in deep inversion:
  High hole concentration at Si-SiO2 interface
   Injection of holes into border traps (~5 nm) dependent on injection time
   Increase of Nox
   VFB shift (depends on time in deep inversion – “waiting time”)
I.Pintilie

## Slide 14: Measurement of Dit
TDRC on MOS capacitor (Thermal Dielectric Relaxation Current)
 | IT(1) | IT(2) | IT(3)
Ec-Eit [eV] | 0.39 | 0.48 | 0.60
FWHM [eV] | 0.26 | 0.13 | 0.07
σeff [cm2] | ~10-15 | ~5∙10-17 | ~10-15
Bias MOS capacitor in strong accumulation  fill interface traps with electrons
Cool to 10 K  freeze traps
Bias to depletion and heat up to 290 K  measure current due to charges released from Si-SiO2 interface
T[K]  Ec-Eit [eV]
I[pA,t]  Dit [(eV∙cm2)-1]
I.Pintilie
J.Zhang
Interpretation not unambiguous !
3 dominant traps:
 Diti vs dose

## Slide 15: GCD*): Principle of Measurement
Measurement of surface current Isurf  Dit (at mid gap)
Isurf
*) GCD = Gate Controlled Diode

## Slide 16: GDC: Surface Current vs. X-ray Dose
I-V curves for irradiated diodes:  Isurface
Results not easy to interpret - Increase up to ~5 MGy
- Decrease beyond 10 MGy
 Saturation expected – Decrease ???
(no Isurf plateau @ high doses, diff. values V↑ and V↓)

## Slide 17: Summary: Results on X-ray Rad. Damage
Parameters used for TCAD Simulations (~ for CIS and Hamamatsu)
  Compare to measurements on segmented sensors
  Optimize AGIPD sensor design for X-ray radiation hardness
Typical values :
Talk by J. Schwandt
TCAD simulations:
 | 0 Gy | 10 kGy | 1 MGy | 10 MGy
Nox [cm-2] | (1-5)∙1011 | 1∙1012 | 3∙1012 | 3∙1012
Isurf [μA/cm2] | 0.005 | 1 | 4 | 6
H.Perrey, J.Zhang
as irradiated
as irradiated

## Slide 18: Impact on p+n Sensors
Sensors irradiated:
  AC coupled from CIS(80 μm pitch)
  DC coupled from Hamamatsu(50 μm pitch)

## Slide 19: Impact on p+n Sensors: Vdepletion
AC coupled CIS sensor :
for pad sensor up to depletion, then C constant
Effects of Nox  increase of electrons in accumulation layer
- step in 1/C2 when undepleted regions below SiO2 separate
- voltage required to deplete entire sensor depends on Nox
 Effect not of relevance for user
J.Zhang
J.Schwandt

## Slide 20: Impact on p+n Sensors: Dark Current
AC coupled CIS sensor:
Interface current (Dit) dominates
- Current from depleted interface (E-field)
- Interface area changes with Vbias
 seen by X-ray users
 minimize depleted interface area
 Important for sensor optimization
J.Zhang
A.Srivastava
Iinterface

## Slide 21: Impact on p+n Pixel Sensors: Breakdown
Simulation 2-dim [x,z and r,z] and 3-dim:
Nox  accumulation layer  changes curvature p+-depletion  changes E-field
Breakdown depends on Nox,tox,p+-implant, Al-overhang, potential on top of sensor (passivation layer), technology, etc.
 Major challenge for sensor design for dose range 0–1 GGy @ > 500 V)
J.Schwandt
z [μ]
z [μ]

## Slide 22: p+n Sensors: Potential on Sensor Surface
Vbreakdown also depends on potential/charge distribution on surface:
- Change of bias voltage  surface fields (non-equilibrium)  movement of surface charges – conductance is a strong function of humidity
- Vbreakdown depends on surface charges
F.G. Hartjes, NIMA 552(2005)168
ATLAS SCT Sensors
(non-irradiated/no Al overhang
Measured inter-strip capacitance vs. time:
humid: ~ 10 min.; dry: days; vacuum: ???
 Effects for sensors with X-ray damage ???
A.Chilingarov et al., NIMA 560(2006)118
humidty:
(T = 22°C)

## Slide 23: p+n Sensors: E-field close to Si-SiO2 interface
For low Nox
500V (equilibrium)  200 V:
 „surplus” of neg. charges on surface
 E-field will point Si  SiO2
 holes will be „lost“ at surface
0V (equilibrium)  200 V:
 „lack” of neg. charges on surface
 E-field will point SiO2Si
 electron „loss“ at surface
For high Nox (X-ray rad. damage)
0V (equilibrium)  200 V:
90% e-loss  30% e-loss
500V (equilibrium)  200 V:
5% e-loss  30% e-loss
Way to measure time con-stants to reach equilibrium
L
R
e-h pairs
----- no losses
- - - - electron losses
------- hole losses
electron losses
hole losses
no losses
T.Pöhlsen
e
600nm laser (λ att ~ 3μm)
h
Signal of R
Al
Al
- - - - - - - - -
- - - - - - - - -
+ + +++++++++++++++++++++++++++++++++++

## Slide 24: Time Constants to reach “Surface Equilibrium”
Results for time constants
- humid: 10’s of minutes
- dry (RH < 5 % @ RT): days
- vacuum: ???
Effect of charge losses:
- only for few μm  ~irrelevant
- charge storage  no pile-up effects as accumulation layers of entire sensor connected
Break-down at high Nox for non-equilibrium: More studies needed !
(in particular for guard ring breakdown in vacuo)
 Efforts needed to optimize properties of “SiO2” and/or passivation layer!
T.Pöhlsen
~ 40 min
~ 80 h

## Slide 25: Summary – Personal Conclusions
X-ray damage is complex:
  Nox and Nit saturate (decrease) for > 10- 100 MGy
   Main parameters measured (understanding? technology !)  Dark current and capacitance not a problem for pixels  Charge losses and pile-up no problem (plasma effect!)
  Main challenge: High breakdown voltage and stable operation
   Proper implementation of parameters into TCAD simulation  Verification simulations with results from segmented sensors  Strategy of sensor optimization  Effects of sensor surface (passivation)  Effects of technology
  What surprises ???
Should XFEL afford building-up expertise in X-ray radiation damage for sensors and electronics ???
Can XFEL afford not

## Slide 26: Thanks
Thanks to:
  Group @ UHH:
    E.Fretwurst
    I.Pintilie (Bukarest + UHH)
    M.Matysek
    T.Pöhlsen
    J.Schwandt
    J.Zhang
  AGIPD @ DESY + PSI + Uni-BN
    + Inspiration from Rainer Richter HLL-München

## Slide 27: Master