Pixel module under X-rays
Alexey Petrukhin, Daniel Pitzl (DESY)
20/04/2012
- X-ray box
- Ag spectrum
- Bias V scan
- Module map
- X-ray test with psi46expert
- Gain calibration for M1207
- Status and plans

---
X-ray box at Uni HH
Uni HH
Bldg 67
Lab 4
2973
35 kV
1 mA
Ag
target
PSI46
test board
pixel module
adapter
front panel operation
-150 V
bias
Cu
target
DAQ via
USB

---
Raw pulse height spectrum
- run 477: 10 s, clock stretch to 25 µs
- tube 35 kV (max), 0.5 mA (50%), Ag target
- 51k trigs, 3.7 M hits, 52 MB (87%)
- all 16 ROCs
- no calibration: gain variations present
- threshold near -250 ADC counts
- Clean above 200 ADC
- Ag peak at -70 ADC

---
Bias voltage scan
- X-ray tube, 35 kV, Ag target
- Full depletion develops between -50 and -80 V bias
100 · clusters / trigger

---
X-ray map through the HDI
bias
filter
cap
power
cable
solder
signal
cable
SMD
caps
SMD
caps
SMD
caps

---
X-ray map through the HDI
bias
filter
cap
power
cable
solder
signal
cable
SMD
caps
SMD
caps
SMD
caps
ROC7
ROC0
TBM
ROC15
ROC8
sensor

---
GUI

---
PSI test procedure (reminder)
- Measure the threshold curves for each ROC :
  - For each value of VthrComp DAC, the fully enabled module is randomly read out several thousand times and the number of hits in each ROC is determined
    - stretch the clock sent to module by up to 65'536 BC to increase the probability of hit finding
  - Fit with error function, set VthrComp and measure Vcal threshold for each pixel. The mean Vcal is taken as a required threshold
                - Fit 3 target measurements with a straight line and get slope (PSI: 65 e/DAC at low Vcal range)
Ba
50%
Ag
Mo

---
DESY x-ray test results
Hit # vs VthrComp fitted by error function
DAC Parameters optimized for lowest threshold
35 kV, 0.03 mA
Ag target
50% point corresponds to the peak pos. of the Ag K-lines and corresponds to 102 Vcal DAC units
A. Petrukhin: X-ray test

---
Gain calibration for M1207
- Calibration is needed for the correct charge reconstruction
- Fit 66560 PH curves (at small Vcal) with a straight line fit: p0, p1
- In data analysis: apply each pair of parameters as calibration factors
- ~3 hours procedure (6 pixels / sec)

---
Gain calibration results
M1207 calibrated Aout, PSI ke conversion
16 ROCs calibrated
x-run 462, Ag target, 35 kV,
0,3 mA, 10 sec.
1 ROC calibrated

---
Status and plans
- Pixel module taken into operation in X-ray box:
  - 50 cm twisted pair cable OK
  - DAC, TB, and TBM parameters determined
  - Threshold trimming applied
- Rate is large, filling test board memory in few seconds:
  - 370 kHz at 10 cm below target (~uniform)
  - 3 MHz at 5 cm below target (ROCs at ends get less rate)
- Silver47 X-rays observed: 22 keV, 6 ke
- Copper29 X-rays below threshold: 8.2 keV, 2.3 ke
- Need other targets: size 5×5 cm2, 0.05‒1.0 mm thickness
  - Niob41 or Molybden42: 16-17 keV, 4.5 ke
  - Tellur52 or Barium56: 27-32 keV (tube limit), 7.5-8.8 ke
  - Order at Goodfellow or American Elements (done by Georg)
- Software installed to uni computer (Ubuntu, 32 bit) – ready to take data
- Target accident: 1 ROC lost (15 are alive)
- Cooling should stabilize analog gain

---
Back up

---
Pixel module
Signal & power
cables
High Density
Interconnect (HDI)
Silicon sensor
16 ROCs
Base stripes

---
psi46 pixel readout chip
4 trim bits
adjustable by programmable DAC, per ROC
programmable register, per pixel

---
psi46 pixel readout chip
adjustable by programmable DAC

---
psi46 DACs
 13 VIBias_Bus 30
 14 Vbias_sf 10
 15 VoffsetOp 55
 16 VIbiasOp 115
 17 VOffsetR0 120
 18 VIon 115
 19 VIbias_PH 130
 20 Ibias_DAC 122
 21 VIbias_roc 220
 22 VIColOr 100
 23 Vnpix 0
 24 VSumCol 0
 25 Vcal 200
 26 CalDel 125
 27 RangeTemp 0
  1 Vdig 6
  2 Vana 150
  3 Vsf 160
  4 Vcomp 10
  5 Vleak_comp 0
  6 VrgPr 0
  7 VwllPr 35
  8 VrgSh 0
  9 VwllSh 35
 10 VhldDel 130
 11 Vtrim 7
 12 VthrComp 124
253 CtrlReg 0
254 WBC 20