## Slide 1: WP-73 TDR ART Review Report: Hard X-Ray Monochromator
On behalf of the development team:
Eng.: X.Dong; Design: M.De Felice, (J.Anton$, S.Kearney$);
PIs: D. Shu$, H.Sinn
$ APS/ Argonne National Lab.
Notes: How to edit the title slide
Upper area: Title of your talk, max. 2 rows of the defined size (55 pt)
Lower area (subtitle): Conference/meeting/workshop, location, date, your name and affiliation, max. 4 rows of the defined size (32 pt)
Change the partner logos or add others in the last row.

## Slide 2: Hard X-Ray beamlines layout
H.Sinn. TDR: X-Ray optics and transport, 2012
Monochromators position
~900 m

## Slide 3: Artificial Channel-Cut Monochromator
Artificial (or pseudo) channel-cut type, non-monolithic
Cryogenic cooling -- Pulse tube cooler
Photo from LCLS/SLAC
Photo from APS/Argonne

## Slide 4: Si-Crystal reflections performances(1)
Intrinsic energy resolution
Diffraction curve
Rocking curve
Si111, Si311, Si511/Si333 are proposed to be used regarding different experimental requirement
Diffraction curve = Darwin curve (one crystal rocking curve);
Rocking curve = two-crystal rocking curve but only one rocking

## Slide 5: Si-Crystal reflections performances(2)
Temperature difference effect on Bragg angle variation
e.g.1st @100K; 2nd @300K
4.5 – 45 deg. Bragg angle to cover required energy range

## Slide 6: Center of rotation (COR)
Option A
Option B
Minimum 1st crystal length, but very long 2nd one;
Increase minimum gap
Almost identical length of both crystals;
Smaller minimum gap

## Slide 7: Geometry proposal
COR
Si(111)
L1=L2=60mm, g=5.5 mm
Si(511)
L1=L2=24mm, g=6 mm
COR
COR@ downstream end edge
extra length as safety factor and extending energy range including Ti-edge

## Slide 8: Motion & initial alignment
Combined Pitch
Pitch
X-tal_1
X-tal_2
Roll

## Slide 9: ACCM-artificial channel-cut mechanisms
Over constrained weak-link structure
(US patent, D.Shu,etc)

## Slide 10: Sine-bar mechanism for combined pitch
f =318.6 Hz
J.Anton, APS

## Slide 11: Steady thermal behavior (Cryo-cooler test)
measurement done by G.Galasso, etc
Heat from 50Ω resistor
PTS8030
3 W

## Slide 12: Dynamic thermal behavior (FEA SImulation)
Calculation done by G.Galasso, etc

## Slide 13: Preliminary design overview
Cone-V-Flat concept
5-DOF alignment to the beam axis
(x, z, pitch, roll, raw)
Z
X
Y
cone
V
flat
Picture from APS

## Slide 14: Prototype project plan
Detailed design after TDR ~2012.10
Standard components (mainly motor/actuator, encoder, controller, etc.) finalizing, ~2012.12
Special parts through APS and test, ~2013.01 Manufacturing drawings, ~2013.01
Tender for production ~2013.03
Delivery, commissioning ~2013.09 +
Purchase, delivery of small series 2014
Questions and comments ? Thank you very much!

## Slide 15: Geometry comparison
 |  |  |  | footprint | 
E (keV) | 4 (mm) | 6 (mm) |  (°) | l0(4) (mm) | l0(6) (mm)
3 | 9.72 | 14.6 | 41.2 | 14.8 | 22.1
5 | 6.28 | 9.42 | 23.3 | 15.9 | 23.9
24 | 1.96 | 2.94 | 4.72 | 23.8 | 35.7
Xtal gap |  |  |  | 2nd Xtal length (appr.) |  |  | 
gmin(4) (mm) |  | gmin(6) (mm) |  | Lmin(4) (mm) |  | Lmin(6) (mm) | 
8.44 | - | 15.41 | - | 112 | - | 198 | -
 | 4.28 |  | 6.41 |  | 62 |  | 92.5
Xtal gap |  |  |  | Xtal length (appr.) |  |  | 
gmin(4) (mm) |  | gmin(6) (mm) |  | Lmin(4) (mm) |  | Lmin(6) (mm) | 
6.46 | - | 9.69 | - | 59 | - | 89 | -
 | 3.42 |  | 5.12 |  | 41 |  | 62
 |  |  |  | footprint | 
E (keV) | 4 (mm) | 6 (mm) |  (°) | l0(4) (mm) | l0(6) (mm)
5 | 6.28 | 9.42 | 49.2 | 8.3 | 12.4
24 | 1.96 | 2.94 | 9.1 | 12.4 | 18.6
Xtal gap |  | Xtal length (appr.) | 
gmin(4) (mm) | gmin(6) (mm) | Lmin(4) (mm) | Lmin(6) (mm)
4.81 | 7.21 | 23.3 | 34.9
Xtal gap |  | 2nd Xtal length(appr.) | 
gmin(4) (mm) | gmin(6) (mm) | Lmin(4) (mm) | Lmin(6) (mm)
6.00 | 9.01 | 42.6 | 64
 |  |  |  | footprint | 
E (keV) | 4 (mm) | 6 (mm) |  (°) | l0(4) (mm) | l0(6) (mm)
8 | 4.63 | 6.95 | 47.9 | 6.24 | 9.36
24 | 1.96 | 2.94 | 14.3 | 7.94 | 11.9
Xtal gap |  | 2nd Xtal length (appr.) | 
gmin(4) (mm) | gmin(6) (mm) | Lmin(4) (mm) | Lmin(6) (mm)
3.92 | 5.88 | 18.9 | 28.4
Xtal gap |  | Xtal length (appr.) | 
gmin(4) (mm) | gmin(6) (mm) | Lmin(4) (mm) | Lmin(6) (mm)
3.45 | 5.18 | 13.1 | 19.6
COR in between x-tals
COR on center of the 1st x-tal
Si111
Si311
Si511

## Slide 16: Exit beam offset & series mono.

## Slide 17: Motion (energy scanning)

## Slide 18: Cooling test set up

## Slide 19: Static thermal effects (FEA simulations)