## Slide 1: The Adaptive Gain Integrating Pixel Detector (AGIPD) Project
Heinz Graafsma, DESY
on behalf of the AGIPD-Consortium

## Slide 2
K. J. Gaffney and H. N. Chapman, Science 8 June 2007
What is the Goal ?

## Slide 3: XFEL Detector requirements
4.5 MHz
What are the Challenges ?

## Slide 4
Single silicon sensor: 128 X 512 PIXELS
readout by 8 x 2 chips; ~ 26 X 105 mm2
Chip carrier board to talk to and readout chips
Hybrid Pixel Array Detectors
Pixellated Silicon Sensor:
absorbs photons and produces electrical signal
Pixellated readout chip (ASIC):
Processes signal produced in sensor and sends data out

## Slide 5
No bulk damage expected for 12 keV (<300 keV)
Surface and interface damage:
  
Higher leakage current
Radiation hardness sensor

## Slide 6
Radiation hardness sensor

## Slide 7
No bulk damage expected for 12 keV (<300 keV)
Surface and interface damage:
  
Higher leakage current
Higher depletion voltage
Lower breakdown voltage
Charge losses at interface
Increased inter-pixel
capacitance
Radiation hardness sensor
Special high voltage design:
withstands 1016 photons
(= 1 GGy )

## Slide 8
Sensors delivered and validated
20 wafers; 40 sensors from SINTEF
Quality control tests:
Flatness
Leakage current
Break-down voltage
Inter-pixel capacitance
Radiation hardness
…
All tests passed!

## Slide 9
The AGIPD RO-Principle
Sensor
Electronics per pixel
ASIC periphery
Chip output driver
Mux
HV
+
-
THR
DAC
SW
CTRL
Analog Mem
Analog Mem
CDS
RO Amp
Calibration circuitry
Adaptive gain amplifier
352 analog memory cells
…
…
RO bus (per column)
Pixel matrix
Notes: Mention random vetoing possibility.

## Slide 10
Dynamic Range

## Slide 11
The Analogue Storage
Sensor
ASIC per pixel
ASIC periphery
Chip output driver
Mux
HV
+
-
THR
DAC
SW
CTRL
Analog Mem
Analog Mem
CDS
RO Amp
Calibration circuitry
Adaptive gain amplifier
352 analog memory cells
…
…
RO bus (per column)
Pixel matrix
Notes: Mention random vetoing possibility.

## Slide 12
Droop curve: Pix 2 cell 128-256

## Slide 13
AGIPD 1.0 Pixel Electronics
200 x 200 micron2 pixels
352 storage cells + veto possibilities.
Minumum signal ~ 300 e- =
  0.1 photon of 12.4keV
Maximum signal ~ 33 106 e- =
  104 photons of 12.4keV
4.5 MHz frame rate
64 x 64 pixels per ASIC
2 x 8 ASICs per module (128x512 pixels, no dead area)
4 modules per quadrant
Chip submitted: 19th April 2013.

## Slide 14
The front-end
Single silicon sensor: 128 X 512 PIXELS
readout by 8 x 2 chips; ~ 26 X 105 mm2
Chip carrier board
to talk to and readout chips
Sensor
Chip

## Slide 15
The module building block
Sensor plus readout ASIC
Analogue board with ADC’s
FPGA + 4 x 10 Ge links

## Slide 16
Folded or seagull design
(Due to limited length of hutch)
Motion system
Closed air loop around electronics
Cooling block
External motors ~ 6W each
Easy to service
Vacuum chamber/flange
can have any shape

## Slide 17
Integration with sample chamber:
Work in progress with EU-XFEL
200mm
SPB needs 100 mm sample detector distance
Closest vacuum flange is > 200 mm away from sample
Working with EU-XFEL engineers to find solution
Likely AGIPD front-end has to stick through gate-valve
Drawings Courtesy
of Eu-XFEL

## Slide 18
The AGIPD RO-Principle
Sensor
Electronics per pixel
ASIC periphery
Chip output driver
Mux
HV
+
-
THR
DAC
SW
CTRL
Analog Mem
Analog Mem
CDS
RO Amp
Calibration circuitry
Adaptive gain amplifier
352 analog memory cells
…
…
RO bus (per column)
Pixel matrix
Notes: Mention random vetoing possibility.

## Slide 19
At the P10 beamline
Beam direction (coming from sample)
AGIPD system
Chip tester box
It took about 1 ½ hours to set up, after about 2 hours we saw the first image
Not in the picture: Sample, Alexanders PC, people, …

## Slide 20
Looking at the direct beam
Gain switching experimentally proven
104 photons / pulse
Single photon sensitivity
4.5 MHz frame rate
photons/second/pixel
10
10
10
10
10
10

## Slide 21
Some XPCS measurements
Scientific quality data obtained
Complete system proven to work
Calibration proven to be adequate
Maxipix
AGIPD
photons/second/pixel
10
10
10
10
10
10

## Slide 22
Sensors delivered; design validated
AGIPD 1.0 submitted 19th April 2013
Interface boards designed and ordered
Single chip assemblies: summer 2013
First modules (8x2 chips): beginning 2014
1k x 1k system(s) beginning 2015
Planning

## Slide 23
Smaller pixels with
  Reduced dynamic range (no gain switching);
  Reduced number of frames?
New Technologies including:
  3D CMOS: second layer (more functionality)
What could be next? AGIPD2?

## Slide 24
3D-CMOS
1) bump-bond to the detector (TSV)
2) amplification & double sampling
3) store the analog signal in a memory cell in 1st/2nd layer (544 memory depth)
4) signal readout
5x5 mm2 double-tier dice
MPW, via CMP
submitted 2011
to be delivered „early“ 2013
Standard Cells

## Slide 25
Smaller pixels with
  Reduced dynamic range (no gain switching);
  Reduced number of frames?
New Technologies including:
  3D CMOS: second layer (more functionality)
  65 nm CMOS technology (more functionality)
  3D CMOS + 65 nm CMOS
What could be next? AGIPD2?
Notes: HGF Portfolio activity

## Slide 26
DESY PSI Uni-HH Uni-Bonn
Julian Becker Roberto Dinapoli Eckhard Fretwurst Markus Gronewald
Laura Bianco Doninic Greiffenberg Robert Klanner Hans Krueger
Peter Goettlicher Beat Henrich Joern Schwandt
Heinz Graafsma Aldo Mozzanica Jiaguo Zhang
Helmut Hirsemann Xintian Shi
Stefanie Jack Bernd Schmitt
Alexandre Klyuev
Sabine Lange
Alessandro Marras
Bjoern Nilsson
Seungyu Rah
Igor Shevyakov
Segrej Smoljanin
Ulrich Trunk
Manfred Zimmer
Who are the People ?
E-XFEL
Jolanta Sztuk-Dambietz