## Slide 1: Status of the XFEL.EU Undulator Systems
J. Pflüger, WP71
Notes: How to edit the title slide
Upper area: Title of your talk, max. 2 rows of the defined size (55 pt)
Lower area (subtitle): Conference/meeting/workshop, location, date, your name and affiliation, max. 4 rows of the defined size (32 pt)

## Slide 2: Overview
XFEL.EU Undulator Systems & Controls
Undulator Segments
  Serial Production of Undulator Segments:
  Representative Results
Intersections
Summary

## Slide 3
Undulator Systems

## Slide 4: European XFEL Overview
Aerial View
SC Accelerator
Dogleg
Switchyard
SASE 1
SASE 2
SASE 3
Dumps
Injector
Accelerator Layout
SASE 1- 3 : Undulator Systems
W. Decking

## Slide 5
e -
e -
SASE2
SASE2 SASE1/3
Downstream View
Tunnel Installation of SASE2 Undulator System – Final WP71 Deliverables
Minor Differences of Components in SASE2 and SASE1/3:
Undulator Segments reversed 180°
Different Arrangement in SASE1 of
  Components on Intersections
  Control Cabinets, cabling
System Sizes:
SASE1: 35 Cells; 0=40mm ; L=213.5m
SASE2: 35 Cells; 0=40mm ; L=213.5m
SASE3: 21 Cells; 0=68mm ; L=128.1m
e -
Intersection 1.1 m
Undulator Segment 5m
Undulator Cell 6.1m
W.Tscheu
Local Control Systems

## Slide 6: Undulator Control System
Developed at XFEL.EU
XFEL.EU guidelines & standards
Based exclusively on industrial Automation Technology
Using the fast EtherCAT protocol
Hardware from Beckhoff GmbH
Running the TwinCAT software
The Topology reflects Hardware
Up to 35 Local Control Systems in Tunnel
Global Control in Computer Service Room
External Access for DOOCS and Experiments
1…. up to 35
S. Karabekyan, A Beckmann, M. Yakopov

## Slide 7
Undulator Segments

## Slide 8: Production of Undulator Segments
The Focus is on Serial Production. R&D is long ago!!
Industrial Part
  Mechanical Support System including Local Motion Control System: 2 Suppliers
  Magnetic Structure
  Documentation in EDMS
XFEL.EU Part
  Integration: Mechanics & Magnetics @ XFEL.EU
  Commissioning @XFEL.EU
  Magnetic Measurement & Tuning
  Preparation for Installation
  Documentation in EDMS

## Slide 9: Production Cycle: Industrial Production
July 2012
Dec 2012
Local Control Systems
Magnetic Structuress
Mechanical Support Systems
Dec 2012
July 2012
Nov 2012
Nov 2012

## Slide 10: Production @ XFEL.EU in Building 36
3 climatized Magnetic Labs
Incoming Magnetic Structures from VAC
Assembly Area for Magnetic Structures
New Undulator Segments
Segments waiting for Magnetic Measurements
Measured and tuned Segments
Hall 36 is rapidly filling up!
Steps @ XFEL.EU:
Mounting of magnetic structures
1st Commissioning
Magnetic Measurements & Tuning
Preparation for Installation
Schedule:
Total Time  2 Years (starting Oct/12)
3 Weeks/Segment
 3 Magnetic Labs needed running in parallel
Scheduled End: Oct / End 2014
December 2012

## Slide 11: Results
Magnetic Measurements
U. Englisch, Y. Li, S.Utermann, F. Wolff-Fabris

## Slide 12: Pole Tuning at XFEL Undulator Segments
Locking Screws
Tuning Screws
 300m
 4mrad
5m
U40: 246 pairs of poles
U68: 144 pairs of poles
BY and BZ tuning is combined
No Shims are used
All errors are tuned by moving Poles
Bipolar Field Corrections

## Slide 13: BY Field: Pole Height Tuning I
FLASH
Determination of ai
Local Field Change
Local gap change
Influence of neighbors
Limit to 3 next neighbors
 300m
Local K-Tuning
Notes: p

## Slide 14: BY Field: Pole Height Tuning II
Tuned Poles:
U40: N =246
U68: N= 144
The Characteristic Coefficients ai were measured for U40/U68:
For the 3 poles next to the end ‘End Poles’
For all other poles ‘Bulk Poles’
Tuning done at the ‘Tuning Gap’: U40: 14mm U68: 16 mm
Most Poles need to be tuned
 One Tuning Step is (usually) sufficient
Bulk poles

## Slide 15: BZ Field: Pole Tilt
 4mrad
BZ is due to errors only.
Ideal BZ trajectory 2nd Field Integral is a Null line
Correction by Tilting Poles
 | 0 [mm] | Tuning Gap [mm] | Tilt Kick Signature [Tmm/mrad]
U40 | 40 | 14 | 0.0127
U68 | 68 | 16 | 0.0325

## Slide 16: Pole Tuning: U68-X003-N001 @ Gap=16mm
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 17: U68-X003 / X004: Comparison of Gap dependent Results
 | Specs | Result
Operational Gap Range [mm] | 10-25 | ---
B0 @ 10mm Gap [T] | 1.69 | 1.66
K @ 10mm Gap | 10.5 | 9.3
RMS I2Y [Tmm2] | < 210 | < 90
RMS I2Z [Tmm2] | < 100 | < 70
Phase Jitter [°] | < 8 | < 8
Transverse Overlap
K-Parameter
Longitudinal Overlap / Phase Jitter

## Slide 18: Pole Tuning : U40-X044-K004 @ Gap=14mm
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 19: U40-X044 / X004: Comparison of Gap dependent Results
 | Specs | Result
Operational Gap Range [mm] | 10-20 | ---
B0 @ 10mm Gap [T] | 1.14 | 1.13
K @ 10mm Gap | 4 | 3.9
RMS I2Y [Tmm2] | < 100 | < 50
RMS I2Z [Tmm2] | < 100 | < 40
Phase Jitter [°] | < 8 | < 7
Transverse Overlap
K-Parameter
Longitudinal Overlap / Phase Jitter
Notes: Before you start editing the slides of your talk change to the Master Slide view: Menu button “View”, Master, Slide Master:
Edit the following 2 items in the 1st slide: 1) 1st row in the violet header: Delete the existent text and write the title of your talk into this text field 2) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
If you want to use more partner logos position them left beside the DESY logo in the footer area Close Master View

## Slide 20: Comparison of Magnetic Results
Measurement Results be independent of the Bench Comparison of Measurements done at all labs
Influence of low temperature exposition
Further Plans: Influence of Transport Conditons to Osdorfer Born
≈2°C

## Slide 21: Present Status of Production
Status of Undulator Production April 19, 2013
Delivery Status:
  43 Support Systems
  38 Magnet structures
Prepared for Measurements: 33
Measured & tuned: 22 Segments, 110m
Waiting for magnetic measurements: 11 Segments
Ready for Installation: 18
In Preparation for Installation: 3
Scheduled End of Production: 12 / 2014

## Slide 22
Intersections

## Slide 23: SASE2 Type Intersection
WP71
Quadrupole Movers
WP71
Phase Shifters
Re-entrant Cavity BPMs+ Electronics
WP17
WP71
Support Bases Adjustors Concrete Pillars
WP71
Air Coil Correctors
Beam Loss Monitors
WP17
WP19
Vacuum System
Status 1/2013:
Delivered: Quadrupoles, WP12
In Production 91 Intersections:
  Phase Shifters (Nov 12), WP71
  BPMS, WP17
  BLMs, WP17
  Vacuum systems, WP19
  Air Coil Correctors, WP71
  Base Plates, WP71
  Adjustors, WP71
  Concrete Pillars , WP71
Released for production:
  Quadrupole Movers (Dec12), Ciemat
Start of Assembly @ XFEL: June 2014
WP12
Quadrupoles

## Slide 24: Phase Shifter Prototypes by IHEP Beijing
Tolerance Band & Operational Range
Period length: 55mm
Gap range: 10.5 to >100mm
Max. Phase Integral @ gap 10.5mm : 23400 T2mm3
1st Hor/Ver Field Integral Tolerance: ±0.004 Tmm
Dynamic Operation without retuning!
2nd Hor/Ver Field Integral Tolerance: ±67 Tmm2
Phase adjustment accuracy: ± 10° (5°)
Min. Gap precision for 5/10° at SASE1/2 ±15/30 µm
H. Lu, S. Sun, C. Wan, IHEP, Beijing
H. Lu, Y. Li, J. Pflueger, NIM A605 (2009),399

## Slide 25: Summary
Production of Undulator Segments in full Swing
  Industrial Production is well in Schedule
  Magnetic measurements & Tuning @XFEL.EU being ramped up
  Magnetic Measurements & Tuning:
    Methods and techniques work fine
    22 Segments tuned to Specs
Intersections:
  Phase Shifter Prototypes in Specs
  (Almost) all Components are in production
Production of Undulator Systems in schedule!

## Slide 26: XFEL.EU Undulator Systems Group, WP71
WP71 Status March 3, 2013, left to right: Yufang Yang, Suren Karabekyan, Xiaoyu Li, Georg Deron, Yuhui Li, Bora Ketenoglu, Uwe Englisch, Joachim Pflüger, Andreas Beckmann, Sonia Utermann, Martin Knoll, Majid Bagha-Shanjani, Fredrik Wolff-Fabris, Hongxian Lin, Karl-Heinz Berndgen, not shown: Mikhail Yakopov

## Slide 27
The End