## Slide 1
Bert Harrop: Diamond Preparation in Princeton
Reactive Ion Etching beneficial for HV stability in sCVD (not in pCVD)
RIE after irradiation could recover signal that was lost during irradiation. (Polarization reduced)
Ideal thickness to remove to be determined (cost-benefit tradeoff), is partial mechanical polish as beneficial?
All new BCM1F sensors should be RIEed
BCML abort diamonds should be RIEed to reduce polarization and recover signal.
Maria Hempel: Diamond measurements in Zeuthen
Characterization of new diamonds well under way.
Optical characterizations for precise thickness and uncovering damage, inclusions etc. (Should we reject bad diamonds?)
Electrical characterization soon with cute new measurement box.
  No packaging needed. Connect with needle.
  Soon automatisation of measurement procedure.
pCVD give trouble lately (change in growth process, or we get the diamonds refused by RD42)

## Slide 2
Erich Grismayer: Characterization for LHC diamonds
Optical inspection to uncover inclusions. Sent back to E6
Electrical characterization: IV and It
  It curve show charge up time constant ~5s
  When break-down currents appear, three time constants: 5s, 100s, 200s
  Could be used to uncover erratic/breakdown limits.
Suggestion to use TCT to measure doping concentration.
  Sensitive to trapped charge as well, bring diamond to de-pumped state?
Konstantin Afanachiev: Characterization of Diamond on Iridium (DOI) samples.
DOI could provide large area detectors, not as pure as sCVD but good homogeneity (no grain boundaries).
  Very potential for pixel detectors.
Partially problems with HV stability.
Measured samples provide CCD slightly better than typical polys.
Need to investigate radiation hardness (polarization)