## Slide 1: First operation of the Gun at XFEL … and the next steps
F.Brinker for the XFEL collaboration
Notes: How to edit the title slide
Upper area: Title of your talk, max. 2 rows of the defined size (55 pt)
Lower area (subtitle): Conference/meeting/workshop, location, date, your name and affiliation, max. 4 rows of the defined size (32 pt)
Change the partner logos or add others in the last row.

## Slide 2
Injector Tunnel ( one year ago )

## Slide 3: First operation of an accelerator at XFEL started December 2013
To make this happen a huge amount of work had to be done
After finalizing the construction work the infrastructure had to be installed in the complete injector building:
Light
Electric power
Cooling water
Air conditioning
IT network
Fire safety
… things I forgot

## Slide 4: We did the gun test during 4 weeks in December and learned a lot about the different systems involved
the operation of the high power RF including modulator, transformer, klystron waveguides
the gun with cathode, solenoids, vacuum system, input coupler and window
the temperature regulation system
the low level RF system
the timing system the machine protection system
the personal interlock system
control system hard- and software

## Slide 5: The location
Magnet power supplies
Klystron, puls- transformer, wave guides
Diagnostics, IT, MPS
Timing, Laser
Second Injector
Gun, Gun cooling
XTIN
XSE
XSIN
UG2
UG3
UG4
UG5
UG6
UG7
Cryo installation

## Slide 6: Personal Interlock
The 2 injectors and the main linac have to be independent
-> Separate regions are needed for the injectors, the main linac, the media shafts and the labyrinths
IL-regions needed for the gun test
Concrete shielding door

## Slide 7: The gun cooling
3 hot water reservoirs

## Slide 8: RF Stations XTIN – UG3
1. Modul Injector #1 Gun
Injector #2 / option
3rd floor / RF stations
6th floor / Inj. #2
7th floor / Inj. #1

## Slide 9
Arrival of the pulse transformers

## Slide 10
and klystrons

## Slide 11
Waveguides – the power is distributed over 4 waveguides and recombined before the gun

## Slide 12: Already getting tight …
From model …

## Slide 13: … to reality
Gun conditioned at DESY-Zeuthen with up to 6MW
Intentionally limited to this value to avoid risk of failures ( RF-window )

## Slide 14: Operational experiences – Thanks a lot for the good preparation !
Magnet power supplies : used for solenoids , OK
Cooling water : OK, minor water temperature issues during the first days at pulse transformer ( power too low )
Modulator : OK beside a single server failure
Pulse transformer, Klystron : OK
Waveguides, circulators directional couplers … : OK
Personal interlock : OK, no failures – operation without temporary access
Timing : OK – some reboots necessary ( unstable power ? )
LLRF : ( no feedbacks used ) no failures but some measurements need further checks
Machine protection : OK, no failures observed
Gun cooling : OK – works very stable under steady state or ramping conditions. Further improvements needed ( and foreseen ) for fast power changes like trips.
Control system : OK – operation completely from BKR ( control room bldg. 30 ), no high level applications yet

## Slide 15: Achieved RF Parameters in December Run
Chart: Forward Power [MW]
| 5.9166666666666666E-2 | 0.58333333333333337 | 1.25 | 2.916666666666667 | 6.4 | 5.9 | 6.15 | 5.9 | 4.5 | 4.6500000000000004 | 4.58 | 3 | 7 | 4.08 | 4.5 | 7.15 | 5 | 4 | 5.0999999999999996 | 5 | 5.2 | 5.2 | 5.4 | 5.4 | 6.3 | 6.25 | 5.4 | 4.8 | 6.56
Dec '14 | 30 | 100 | 300 | 40 | 200 | 100 | 100 | 410 | 325 | 355 | 650 | 55 | 420 | 600 | 50 | 650 | 650 | 650 | 650 | 300 | 650 | 550 | 650 | 400 | 400 | 650 | 650 | 300 | 650
specification | 650
RF Run from 9th to 23rd of Dec.
Gun has been conditioned at Zeuthen to 6MW – no “Gun” interlocks observed ( no vacuum break out or reflected power ILs )

## Slide 16: Development of pulse energy over time
Chart: date
| 9.12. | 10.12. | 11.12. | 12.12. | 13.12. | 14.12. | 14.12. | 15.12. | 16.12. | 16.12. | 17.12. | 17.12. | 17.12. | 17.12. | 18.12. | 18.12. | 19.12. | 19.12. | 19.12. | 20.12. | 20.12. | 21.12. | 21.12. | 22.12. | 23.12. | 23.12. | 23.12. | 23.12. | 23.12.
Pulse Energy [J] | 1.775 | 58.3333 | 375 | 0 | 256 | 1180 | 615 | 590 | 1845 | 1511.25 | 1625.9 | 1950 | 385 | 1713.6 | 2700 | 357.5 | 3250 | 2600 | 3315 | 3250 | 1560 | 3380 | 2970 | 3510 | 2520 | 2500 | 3510 | 3120 | 1968
Goal :
6.5MW ∙ 650µs = 4.2 kJ

## Slide 17: So, what limited the power / pulse length ? → Interlocks only from light sensors at RF window
Signal from photodiode on air side ( Spark ) increases slowly with pulse energy
Photomultipliers on vacuum side detect only high spikes together with spark det.
One exception only after 6:10, Dec 23rd for about 1h with light on vacuum side

## Slide 18: Hint to problems with RF spring:
Strong vacuum activity behind the cathode – no activity seen within the gun
Same behavior when changing the magnetic field. Ramping up the solenoids with RF-power without any problems – beside pressure increase in the cathode chamber
No interlocks !
After the test the gun was opened to examine the cathode holder
The holder is a new design after damages have been observed at the gun backplane at FLASH

## Slide 19: Massive damages of Cathode spring after operation

## Slide 20: Improvement seen at the next gun at Zeuthen
The gun 4.4 has been opened after operation at Zeuthen last week
Cathode spring in much better condition
No melting
Nevertheless 3 broken fingers

## Slide 21: Plan for improvements:
Improvement of surfaces (test in Zeuthen starts now ):
  Cathode holder have been gold plated
  Spring has been rhodium and gold coated
  Cathode plugs have been electro polished
In parallel : modification of the geometry
  A different spring geometry ( “watchband design” used at FLASH ) is in preparation for this autumn
Cathode plug
backplane

## Slide 22: Attacking the window problem
The light emission seen at the window is a hint for RF-induced discharges
These can be reduced by ultra high vacuum in the waveguides, fillings of SF6 or increasing the air pressure
  SF6 is practically excluded due to safety reasons
  For UHV operation waveguide components like in-vacuum T-combiner and directional couplers are needed – these are presently under tests at Zeuthen
  At XFEL the waveguides will be prepared for higher air pressure of up to 3 bar for the next run

## Slide 23
Further installations until mid ‘14
Add dispersive arm and diagnostics to the gun
Major part of the warm beamline
Injector dump
Cryo installations for injector
Laser system

## Slide 24
Installation of Cryo transfer lines

## Slide 25
Installation of Cryo-Module supports

## Slide 26: next : Gun Section
Charge
  Toroid
  Dark current monitor
  Faraday Cup
  Integrated Faraday Cup
Position
  BPM
Beam shape
  Screens
Energy
  Spectrometer magnet
  Screen
Notes: Before you start editing the slides of your talk change to the Master Slide view:
Menu button “View” > Slide Master:
Edit the following items:
1. On the Title slide (second master slide)
a) Click to add title of your talk
b) Click to add subtitle (conference, location, name of the speaker, date)
2. On the Slide Master (first master slide)
a) The 1st row in the violet header: Delete the existent text and write the title of your talk into this text field b) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
Close Master View

## Slide 27: UV-Laser (MBI) installation in UG5 started, beamline design ready
(FS-LA)

## Slide 28: Diagnostics : Part 1
Charge
  Toroid:
  DCM:
Position
  BPMC/R:
  BPMA:
  BPMF:
Beam Shape
  OTRL:
    Mover + Screen:
    Chamber: to be delivered by Uni Upsala
  Loss monitors:
    BLM:
TDS will come later in 2014
Notes: Before you start editing the slides of your talk change to the Master Slide view:
Menu button “View” > Slide Master:
Edit the following items:
1. On the Title slide (second master slide)
a) Click to add title of your talk
b) Click to add subtitle (conference, location, name of the speaker, date)
2. On the Slide Master (first master slide)
a) The 1st row in the violet header: Delete the existent text and write the title of your talk into this text field b) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
Close Master View

## Slide 29: Diagnostics Part 2
Charge
  Toroid:
Position
  BPMA:
  BPMD:
Beamshape
  OTRC:
  OTRD:
Losses
  Loss monitors
Fast Kicker allow the selection of bunches for diagnostics
Kontrolle des Strahls am Dump
Notes: Before you start editing the slides of your talk change to the Master Slide view:
Menu button “View” > Slide Master:
Edit the following items:
1. On the Title slide (second master slide)
a) Click to add title of your talk
b) Click to add subtitle (conference, location, name of the speaker, date)
2. On the Slide Master (first master slide)
a) The 1st row in the violet header: Delete the existent text and write the title of your talk into this text field b) The 2 rows in the footer area: Delete the text and write the information regarding your talk (same as on the Title Slide) into this text field.
Close Master View

## Slide 30: Laser Heater system with undulator and infrared Laser
OS0
  Telescope
  photon waist size and position
  4D laser stabilisation system
  2 Attenuation systems
OS1
Diagnostics after undulator
Online Monitoring
  Power
  Position Sensitive Det. (PSD)
  photon beam size (camera)
  Timing (photo diodes)

## Slide 31: Magnets are in house and RFI Girder pre-installation started

## Slide 32: Dump vehicles ready and tested

## Slide 33: Schedule for the next steps

## Slide 34: Longer perspective until mid of 2015

## Slide 35
Thank You !