## Slide 1: Double-Sided Segmented Solid State Detectors
Position-, Energy-, and Timesensitive Detectors applied to X-ray Imaging and Compton Polarimetry
Karl-Heinz Blumenhagen (HI-Jena)
Ge(i) Detector for X-ray Imaging and Compton Polarimetry
56 mm
32 mm
11 mm
p+ contact (128 strips)
a-Ge contact (48 strips)
Si(Li) Detector for Compton Polarimetry
7 mm
64 mm
64 mm
p+ contact (32 strips)
Li doped contact (32 strips)
Spatial intensity distributions for different polar scattering angles at ħ = 214.8 keV.
Characterization of the Si(Li) Compton polarimeter at PETRA III (DESY)
Current development: Si(Li)- Polarimeter with FEBEX2-based DSP-electronics to replace classical electronics. FEBEX2 features:
2D-Strip Detectors in the Cystal Spectrometer FOCAL
Digital Readout
(In the framework of HDRI)
Energy spectrum of the incident synchrotron radiation.
Digitized data to PC
FEBEX2
FEBEX2
...
8 analog
inputs per FEBEX2
FEBEX2
65 Msamples/s, 12 bit resolution for each channel
2.5 Gbit/s data transfer rate
Trigger and time stamp logic with onboard FPGA
Easily scalable due to “chaining“
“chaining“
Crystal: n-type Germanium, thickness: 11 mm
Segmentation: 48 strips (1167 µm wide) x 128 strips (220 µm wide)
Crystal: p-type Silicon, thickness: 7 mm
Segmentation: 32 strips (2 mm wide) x 32 strips
(2 mm wide)
Dedicated detector test beamtime
Highly polarized hard x-rays available
Goal: determination of polarimeter quality
counts
3rd
4th
2nd
harmonic
Energy [keV]
Principle: energy is measured via the position using Bragg reflection.
2012 run: setup and preliminary results.
2006 test run: 2D-detector images with:
time condition (top)
time and energy condition (bottom)
with time condition
with time and energy condition