## Slide 1: Dependence of cavity performance on the cooling dynamics around Tc
Alexander Romanenko, Anna Grassellino, Dmitri Sergatskov, Olexandr Melnychuk
TTC, DESY
24th March 2014

## Slide 2: Preamble
Colleagues at HZB had found that for dressed BCP cavities the details of cavity cooling can impact significantly residual resistance
  Suggested explanation was that thermocurrents generated in dressed cavities produce additional magnetic field, which gets trapped
  Also suggested that slower cooling through Tc is better from samples studies
Cornell had implemented slow cooling procedure in HTC and measured a low Rres on a 120C bake + HF rinse cavity
Final recommendation at SRF13 for minimization of residual resistance was very slow uniform cooling through Tc
We decided at FNAL to start investigating the findings on bare cavities in VTS

## Slide 3: Cooling details
Slow cooling: <0.3K/min
Fast cooling: 1.8-2.4K/min
Four temperature sensors were positioned differently in the experiments, example for one 9-cell test
Generally one T sensor next to each fluxgate probe

## Slide 4: First studies at FNAL in VTS – nitrogen doped bare 9-cell
Slow cooling lead to strong deterioration of Q0 in VTS
Oct, 2013
Nitrogen doping method described in :
A Grassellino et al 2013 Supercond. Sci. Technol. 26 102001

## Slide 5: Deterioration solely due to residual resistance
Increase in residual resistance at low field plus a field dependent component (as expected for trapped flux)
Decomposition method described in :
A. Romanenko, A. Grassellino Appl. Phys. Lett. 102 , 252603 (2013)

## Slide 6: Performed extensive systematic studies since then
Used multiple cryogenic fluxgate probes attached to the cavity walls to directly see what magnetic field environment is during cryocycles
  Most importantly what happens around Tc (9.25K)
    Directly observe the Meissner effect
Explored all state-of-the-art treatments (EP, EP+120C, N doping), fine and single grain material
  Tens of tests on 3 nine-cell cavities and 8 single-cell cavities
The same underlying mechanism always observed
  Slow cooling leads to much LOWER external flux expulsion (more trapping) than fast cooling
    Hence the difference in residual resistance
  Magnetic field before transition is similar for fast/slow cooling => no effect of thermal currents seen on bare cavities
See http://arxiv.org/abs/1401.7747 for the preprint

## Slide 7: Degradation of Q after slow cooling observed for all surface treatments
LG EP
FG EP
FG EP+120C bake
FG EP N doped

## Slide 8: Degradation of residual resistance after slow cooling observed for all surface treatments (simultaneous cooling)
LG EP
FG EP

## Slide 9: Understanding why: magnetic probe positioning - examples
H=5mG
Field enhancement
due to Meissner effect – ambient flux expulsion
Fluxgate
magnetometers
mG

## Slide 10: What did magnetic probes reveal?
Similar magnetic field at transition between fast/slow
  No additional field generated during fast cooling
Systematic strong difference in flux expulsion efficiency between slow and fast cooling
EP+120C,
Fine grain
N doped,
Fine grain
EP, fine
grain

## Slide 11: Examples: Single and Nine cell EP + 120C bake

## Slide 12: Expulsion efficiency after fast cooling from different temperatures
Same effective expulsion for different starting temperatures (=different thermal gradients) for the fast cooldown – OK even from 11K

## Slide 13: One of the possible mechanisms (proposed by D. Sergatskov, TBP)
  Fast cooldown – well-defined superconducting/normal boundary is moving from bottom of the cavity to the top – no energy barrier for flux to be expelled
Superconducting
(T<Tc)
Normal conducting
(T>Tc)

## Slide 14: Slow uniform cooldown
  Slow uniform cooldown – superconductivity is nucleated at multiple spots which reach T<Tc
  Flux surrounded by superconducting areas has an energy barrier for escaping=> more trapping is possible

## Slide 15: Are these findings in contradiction with previous findings?
Even if conclusions are opposite, not necessarily, perhaps only at first glance. (careful when comparing results, evil in the details):
Effect is found on bare cavities – so while studies cannot say if additional effects come from dressing (eg thermocurrents), the effect we discovered will always be present, dressed or undressed
Underlying mechanism is same (almost no flux expelled with slow uniform cooling), however trapped flux translates differently into residual resistance for different surface processing: size of vortex cores, normal state resistivity and pinning strength depend on surface mfp, which is eg 2 nm for 120C bake, 40 nm for nitrogen doping, >500 nm for EP/BCP – so effect will manifest with different magnitude for different surface processing and field levels

## Slide 16: Are these findings in contradiction with previous findings?
Even if conclusions are opposite, not necessarily, perhaps only at first glance. (careful when comparing results, evil in the details):
Magnitude of the effect will depend also on level of shielding/ambient field level which will vary from dewar to dewar
4. Definition of “fast” and “slow”:
  Rate is not determined by K/hr but by the instantaneous cooling rate through Tc which will vary upon slow cooling procedure
  Faster than what? Slower than what? Appears to be a threshold effect
  Homogeneous vs not

## Slide 17: Possible alternative interpretation of HZB results
Faster cooldown leads to lower residual resistance -> fully consistent with Fermilab results on bare cavities
  This possible interpretation is also acknowledged in the HZB paper
J.-M. Vogt, O. Kugeler, and J. Knobloch, Phys. Rev. ST Accel. Beams 16, 102002 (2013)

## Slide 18: Cornell Data - HTC better than VTS on the same cavity
Possible “contradiction” #1: HTC (very slow cooling) gets better Q0 than VTS
  But VTS cooling was also intentionally slow : “..the cavity was warmed up to above 10 K, and slowly cooled through the critical temperature at a rate of ∼ 1 K/hr” and with much higher ambient field
=> Our findings explain lower Q0 (higher Rres) in VTS vs HTC, and also subpar VTS Q0 of ERL 7.1 cavity
N. Valles, PhD thesis, Cornell University, 2014
N. Valles et al, SRF’13

## Slide 19: Cornell Data - Thermal cycling
Possible “contradiction” #2: Warming up above Tc and cooling back down slowly leads to better Q0
  But improvement over initial VERY SLOW cooldown
    Also have to compare instantaneous dT/dt (not average) to define which is slower
=> Does not contradict our findings
N. Valles, PhD thesis, Cornell University, 2014
Also notice NO decrease of Q0 after fast cooldown
N. Valles et al, SRF’13

## Slide 20: Conclusions
Strong effect of cooling rate through transition T on flux trapping/expulsion
Reflects differently on losses depending on surface processing
  understandable as size of vortex cores, normal state resistivity and pinning strength depend on surface mfp
Recommendation for cooldown in cryomodules: cool with rate dT/dt >~ 30 mK/sec (1.8K/min, could be lower, TBD)
  Starting T does not appear to matter but to be further investigated in dressed cavities
  Results from DESY (J. Sekutowicz) on full cryomodule with large grain EP+120C baked cavities confirm this recommendation
Next we will repeat studies on dressed cavities to study if additional mechanisms due to the dressing occur or not (perhaps already well known?…see next slide)

## Slide 21: DESY dressed cavities data analysis by Alex Melnychuk

## Slide 22: Acknowledgements
Processing: Allan Rowe and his FNAL/ANL SRF processing team, Curtis Crawford
VTS cryo-capabilities: Bruce Squires and his FNAL cryo group, Yuri Pischalnikov